## ACO2
Sampo\D:\aco2\data\SMYYMMDD.aco
Datafile column | Nokeval5 Channel | Measured Quantity | Unit in datafile | Notes
1 |  | date and time | yyyymmddhhmmss
2 | 1 | H2O URAS [g/m3] | mV
3 | 2 | CO2 URAS [ppm] | mV
4 | 3 | H2O URAS pressure [hPa] | mV
5 | 4 | CO2 URAS pressure [hPa] | mV
6 | 5 | MFC URAS [Nml/min] | mV
7 | 6 | Empty
8 | 7 | Empty
9 | 8 | URAS Temperature [°C] | mV
10 | 13 | CO2 LiCor [ppm] | mV
11 | 14 | H2O LiCor [ppth] | mV
12 | 15 | LiCor Temperature [°C] | mV
13 | 16 | LiCor pressure [hPa] | mV

## Aqua
Sampo\D:\smear\Aqua\SMYYMMDD.AQU |  |  |  | Began 7.6.1997
Column | Com | Channel | Meaning | Unit
| Sampo 5 | Promux |  |  |  |  | Conversion
1 |  |  | date and time | yyyymmddhhmmss
2 |  | 1 | Rain 0 | counts | The one on the ground | ARG 100 | rmm = (0.1990*count)mm
3 |  | 2 | Rain 1 | counts | PIT 50 | Rainer 1 | rmm = (0.0849*count)mm
4 |  | 3 | Rain 2 | counts | PIT 70 | Rainer 2 | rmm = (0.0788*count)mm
5 |  | 4 | Rain 3 | counts | PIT 100 | Rainer 3 | rmm = (0.0752*count)mm
6 |  | 5 | Rain 4 | counts | PIT 130 | Rainer 4 | rmm = (0.0704*count)mm
7 |  | 6 | Rain 5 | counts | PIT 160 | Rainer 5 | rmm = (0.0801*count)mm
8 |  | 7 | Rain 6 | counts | PIT 80 | Rainer 6 | rmm = (0.0817*count)mm
9 |  | 8 | Rain 7 | counts | PIT 60 | Rainer 7 | rmm = (0.0800*count)mm
10 |  | 9 | D1 soil outflow | counts |  |  | flow = 1.0 dm3/count
11 |  | 10 | D1 surface outflow | counts |  |  | flow = 1.0 dm3/count
12 |  | 11 | D2 soil outflow | counts |  |  | flow = 1.0 dm3/count
13 |  | 12 | D2 surface outflow | counts |  |  | flow = 1.0 dm3/count
14 |  | 13 | empty |  | output 0
15 |  | 14 | empty |  | output 0
16 |  | 15 | empty |  | output 0
17 |  | 16 | empty |  | output 0

## Dynamax
Sampo\D:\smear\Dynamax\SMYYMMDD.DYN |  |  | Began 15.7.1997
Column | Label | Meaning
1 | I.D | array ID
2 | day | number of the day
3 | time | measuring time | min
4 | Flo_G1 | flow in gauge 1 (G1) | g/h
5 | Ksh_G1 | dT_G1 | thermal conductivity of the stem * temperature increase of sap in G1 | W/m*C
6 | °C
7 | Pin_G1 | power in to the stem in G1 | W
8 | Qv_G1 | vertical heat flux in G1 | W
9 | Qr_G1 | radial heat flux in G1 | W
10 | Qf_G1 | sap flow heat flux in G1 | W
11 | Facc_G1 | total flow in G1 | g
12 | Flo_G2 etc. to Facc_G8 | following measured values from gauges 2 - 8 (G2 - G8)
Sampo\D:\smear\Dynamax\SMYYMMDD.DYN |  |  | Began 20.6.2007
Column | Meaning | Unit | Notes
1 | array ID
2 | Year
3 | Julian day
4 | Time | hour:min
5 | Empty
6 | Temperature sensor | °C
7 | Temperature sensor | °C
8 | Empty
9 | Empty
10 | Temperature sensor | °C
11 | Temperature sensor | °C
12 | Heated sensor | °C
13 | Heated sensor | °C
14 | Temperature sensor | °C
15 | Temperature sensor | °C
16 | Heated sensor | °C
17 | Heated sensor | °C
18 | CO2 | ppm
19 | Empty
20 | CO2 | ppm
21 | CO2 | ppm
22 | CO2 | ppm
23 | CO2 | ppm
Sampo\D:\smear\Dynamax\SMYYMMDD.DYN |  |  | Began 4.3.2008, updated 20.5.2010
Column | Meaning | Unit | Notes
1 | array ID
2 | Year
3 | Julian day
4 | Time | hour:min
5 | Empty
6 | Temperature sensor | °C | reference temperature for tree 1? (cable 2)
7 | Temperature sensor | °C | reference temperature for tree 3? (cable 3)
8 | Empty |  | reference temperature ?
9 | Empty |  | reference temperature ?
10 | Temperature sensor | °C | Tree 1 cable 2 tip of needle
11 | Temperature sensor | °C | Tree 3 cable 3 tip of needle
12 | Heated sensor | °C | Tree 1 cable 2 base of needle
13 | Heated sensor | °C | Tree 3 cable 3 base of needle
14 | Temperature sensor | °C | Tree 3 tip?
15 | Temperature sensor | °C | Tree 3 base?
16 | Heated sensor | °C | Tree 3 tip?
17 | Heated sensor | °C | Tree 3 base?
18 | Steady-temp sapflow |  | "Tree 4", 5m from the tower, cable 5
19 | Granier sapflow | mV | "Tree 4", 5m from the tower, to get °C multiply with ~30, cable 6
20 | Steady-temp sapflow |  | Tree 1, cable 1
21 | empty
22 | empty
23 | empty
24 | Granier sapflow | mV | Tree 2, cable 4
Sampo\D:\smear\Dynamax\SMYYMMDD.DYN |  |  | Began 20.5.2010
Column | Meaning | Unit | Notes
1 | array ID
2 | Year
3 | Julian day
4 | Time | hour:min
5 | Empty
6 | Empty
7 | Empty
8 | U diff
9 | Empty
10 | Empty
11 | Empty
12 | Empty
13 | Empty
14 | Empty
15 | U diff
16 | U diff
17 | U diff
18 | U power
19 | U power
20 | Empty
21 | Empty
22 | Empty
23 | U diff
24 | U diff
Sampo\D:\smear\Dynamax\SMYYMMDD.DYN |  |  | Began 1.7.2010
Column | Meaning | Unit | Notes
1 | array ID
2 | Year
3 | Julian day
4 | Time | hour:min
5 | Empty
6 | Empty
7 | U difference | uV | Tree 5 / Cambpell Channel 2/2
8 | U difference | uV | Tree 3 / Campbell Channel 6/1
9 | U difference | uV | Tree 1 upper / Campbell Channel 6/2
10 | U power | mV | Tree 4 south / Campbell Channel 7/1
11 | U difference | uV | Tree 4 north / Campbell Channel 7/2
12 | U power | mV | Tree 1 lower / Campbell Channel 8/1
13 | Empty
14 | U difference | uV | Tree 2 lower / Campbell Channel 10/2
15 | U difference | uV | Tree 2 upper / Campbell Channel 10/1
Tree map location: \\Timo\smear\dutyinfo\TreeMap/sapflow_map_2010.pdf

## Dendrometers
Began 29.4.2010
Pepe\D:\smear\DendroNok\SMYYMMDD.dnd
Datafile column | Nokeval0 Channel | Measured Quantity | Units in datafile | Notes
1 |  | date and time | yyyymmddhhmmss
2 | 1 | D 122 NE45
3 | 2 | D 127 NE45
4 | 3 | D 171 NE45
5 | 4 | D 188 SE45
6 | 5 | D 190 SE45
7 | 6 | D 179 E50
8 | 7 | D 183 E50
9 | 8 | Empty
10 | 9 | Empty
11 | 10 | Empty
12 | 11 | Empty
13 | 12 | Empty
14 | 13 | Empty
15 | 14 | Empty
16 | 15 | Empty
17 | 16 | Box temperature | °C

## Equitensiometers
Equitensiometers dug in the soil during the first week of June -05. | [5 empty] | New sensors were installed 22.04.2008
Equitensiometer wires installed 17.6.05 | [5 empty] | Equitensiometer wires installed 17.6.05
PhotoNokeval M04 | [5 empty] | PhotoNokeval M04
Louhi\D:\smear\eqtensio\SMYYMMDD.eqi | [5 empty] | Louhi\D:\smear\eqtensio\SMYYMMDD.eqi
Datafile column | Nokeval 4 Channel | Measurement | Units of measurement | Descriptions | Notes | Datafile column | Nokeval 4 Channel | Measurement | Units of measurement | Descriptions | Notes
1 |  | date and time | yyyymmddhhmmss |  |  | 1 |  | date and time | yyyymmddhhmmss
2 | 1 | Eq2 #021 | mV |  |  | 2 | 1 | Eq2 #021 | mV
3 | 2 | Eq2 #022 | mV |  |  | 3 | 2 | Eq2 #022 | mV
4 | 3 | Eq2 #023 | mV |  |  | 4 | 3 | Eq2 #023 | mV
5 | 4 | Eq2 #025 | mV |  |  | 5 | 4 | Eq2 #025 | mV
6 | 5 | Eq2 #026 | mV |  |  | 6 | 5 | Eq2 #026 | mV
7 | 6 | Eq2 #028 | mV |  |  | 7 | 6 | Eq2 #028 | mV
8 | 7 | Eq2 #029 | mV |  |  | 8 | 7 | Eq2 #029 | mV
9 | 8 | Eq2 #027 | mV |  |  | 9 | 8 | Eq2 #027 | mV
10 | 9 | Eq2 #030 | mV |  |  | 10 | 9 | Eq2 #030 | mV
11 | 10 | empty |  |  |  | 11 | 10 | Growth sensor | mV
12 | 11 | empty |  |  |  | 12 | 11 | Empty
13 | 12 | empty |  |  |  | 13 | 12 | Empty
14 | 13 | empty |  |  |  | 14 | 13 | Temp. sensor | °C
15 | 14 | empty |  |  |  | 15 | 14 | Temp. sensor | °C
16 | 15 | empty |  |  |  | 16 | 15 | Temp. sensor | °C
17 | 16 | Net rad / Reemann | mV |  | kampanjakäyttö | 17 | 16 | Temp. sensor | °C

## Gaslog
Eddyold\D:\gaslog\SMYYMMDD.pet
Began 30.08.2007
Datafile column | Nokeval 0 Channel | Measurement | Units of measurement
1 |  | date and time | yyyymmddhhmmss
2 | 1 | Horiba CO analyser inlet dp[hPa] | hPa
3 | 3 | Mast 16.8 m sample line dp [hPa] | hPa
4 |  | empty
5 |  | Cabin temperature | C
6 |  | cal. Instrument
7
8
9
10

## CanPAR
Began 24.06.2010
REA\C:\canpar\AEYYMMDD.cp1
Datafile column | Measured Quantity | Unit in datafile | Notes
1 | date and time | yymmddhhmmss
2 | -
3 | Site address | 1,2,3 | 1=beside REA cabin, 2=beside path, 3=behind Photo Tower
4 | 1
5 | Logan PAR | mV
6 | Logan PAR | mV
7 | Logan PAR | mV
8 | Logan PAR | mV
9 | Logan PAR | mV
10 | Bat. Voltage | mV

## Maapar
During the measurement campaign 26.6.-8.7.05. |  |  | Began 31.10.2003 |  | Nokeval configuration modified 22.4.2008.
Pepe\D:\smear\maaparnor\SMYYMMDD.mpr
Datafile column | Nokeval5 Channel | Measured Quantity | Unit in datafile | Notes
1 |  | date and time | yyyymmddhhmmss
2 | 1 | Licor PAR | mV
3 | 2 | Licor PAR | mV
4 | 3 | Licor PAR | mV
5 | 4 | Licor PAR | mV
6 | 5 | Net rad Reemann | mV
7 | 6 | Net rad Reemann | mV
8 | 7 | Net Reemann | mV | Empty since 28.10.2005
9 | 8 | Net Rebs # 5 | mV | Between 03.06.2005 - 28.10.2005
10 | 9 | Net Rebs # 4 | mV | Between 03.06.2005 - 21.09.2005
11 | 10 | Net Rebs "Sweden" | mV | Between 23.06.2005 - 21.09.2005
12 | 11 | Vaisala #1750002, 15 cm | mV | Between 23.06.2005 - 21.09.2005
13 | 12 | Vaisala # 1720008 50 cm | mV | Between 23.06.2005 - 21.09.2005
14 | 13 | Vaisala # 2010003 150 cm | mV | Between 23.06.2005 - 21.09.2005
15 | 14 | CO2 Licor 840 | mV | Between 28.06.2005 - 21.09.2005
16 | 15 | H20 Licor 840 | mV | Between 28.06.2005 - 21.09.2005
17 | 16 | Temperature in REA cottage | °C | Began 12.06.2005
The measurement range of Vaisala was 0..1000ppm
Output voltage 0..2500 mV
"Vaimennus/damping" 1/5
Nokeval 0..500 mV equals to 0..1000 ppm.
Pepe\D:\smear\maaparnor\SMYYMMDD.mpr
Datafile column | Nokeval5 Channel | Measured Quantity | Unit in datafile | Notes
1 |  | date and time | yyyymmddhhmmss
2 | 1 | Licor PAR | mV
3 | 2 | Licor PAR | mV
4 | 3 | Licor PAR | mV
5 | 4 | Licor PAR | mV
6 | 5 | Net rad Reemann | mV
7 | 6 | Net rad Reemann | mV
8 | 7 | 20mV temp.profile | mV
9 | 8 | 20mV temp.profile | mV
10 | 9 | 20mV temp.profile | mV
11 | 10 | 20mV temp.profile | mV
12 | 11 | temp.profile reference | °C
13 | 12 | VOC pit160 soil temp. | °C | Since 23.04.2008
14 | 13 | VOC pit160 soil temp. | °C | Since 23.04.2008
15 | 14 | Theta probe | mV | Since 01.04.2009
16 | 15 | Theta probe | mV | Since 01.04.2009
17 | 16 | Actinic Flux filterradiometer | °C | Between 10.06.2009 - 10.09.2009

## TDR
Sampo\D:\smear\TDR\SMyymmdd.TDR |  | 9.5.1997 - 8.12.2004
Updated 8.9.2004
column | label | type of measurement | column | label | type of measurement
1 | ? | time | 1 | ? | time
2 | Year | time | 2 | Year | time
3 | Day of Year | time | 3 | Day of Year | time
4 | Time | time | 4 | Time | time
5 | ? | ? | 5 | ? | ?
6 | 50/H/Water | volumetric water concentration | 6 | 50/H/Water | volumetric water concentration
7 | 50/A/Water | volumetric water concentration | 7 | 50/A/Water | volumetric water concentration
8 | 50/B/Water | volumetric water concentration | 8 | 50/B/Water | volumetric water concentration
9 | 50/C/Water | volumetric water concentration | 9 | 50/C/Water | volumetric water concentration
10 | 50/H/Electric | electric conductivity | 10 | 50/H/Electric | electric conductivity
11 | 50/A/Electric | electric conductivity | 11 | 50/A/Electric | electric conductivity
12 | 50/B/Electric | electric conductivity | 12 | 50/B/Electric | electric conductivity
13 | 50/C/Electric | electric conductivity | 13 | 50/C/Electric | electric conductivity
14 | 70/H/Water | volumetric water concentration | 14 | 70/H/Water | volumetric water concentration
15 | 70/A/Water | volumetric water concentration | 15 | 70/A/Water | volumetric water concentration
16 | 70/B1/Water | volumetric water concentration | 16 | 70/B1/Water | volumetric water concentration
17 | 70/B2/Water | volumetric water concentration | 17 | 70/B2/Water | volumetric water concentration
18 | 70/C/Water | volumetric water concentration | 18 | 70/C/Water | volumetric water concentration
19 | 70/H/Electric | electric conductivity | 19 | 70/H/Electric | electric conductivity
20 | 70/A/Electric | electric conductivity | 20 | 70/A/Electric | electric conductivity
21 | 70/B1/Electric | electric conductivity | 21 | 70/B1/Electric | electric conductivity
22 | 70/B2/Electric | electric conductivity | 22 | 70/B2/Electric | electric conductivity
23 | 70/C/Electric | electric conductivity | 23 | 70/C/Electric | electric conductivity
24 | 100/H/Water | volumetric water concentration | 24 | 100/H/Water | volumetric water concentration
25 | 100/A/Water | volumetric water concentration | 25 | 100/A/Water | volumetric water concentration
26 | 100/B1/Water | volumetric water concentration | 26 | 100/B1/Water | volumetric water concentration
27 | 100/B2/Water | volumetric water concentration | 27 | 100/B2/Water | volumetric water concentration
28 | 100/C/Water | volumetric water concentration | 28 | 100/C/Water | volumetric water concentration
29 | 100/C2/Water | volumetric water concentration | 29 | 100/C2/Water | volumetric water concentration
30 | 100/C3/Water | volumetric water concentration | 30 | 100/H/Electric | electric conductivity
31 | 100/H/Electric | electric conductivity | 31 | 100/A/Electric | electric conductivity
32 | 100/A/Electric | electric conductivity | 32 | 100/B1/Electric | electric conductivity
33 | 100/B1/Electric | electric conductivity | 33 | 100/B2/Electric | electric conductivity
34 | 100/B2/Electric | electric conductivity | 34 | 100/C1/Electric | electric conductivity
35 | 100/C1/Electric | electric conductivity | 35 | 100/C2/Electric | electric conductivity
36 | 100/C2/Electric | electric conductivity | 36 | 130/H/Water | volumetric water concentration
37 | 100/C3/Electric | electric conductivity | 37 | 130/A/Water | volumetric water concentration
38 | 130/H/Water | volumetric water concentration | 38 | 130/B1/Water | volumetric water concentration
39 | 130/A/Water | volumetric water concentration | 39 | 130/B2/Water | volumetric water concentration
40 | 130/B1/Water | volumetric water concentration | 40 | 130/C1/Water | volumetric water concentration
41 | 130/B2/Water | volumetric water concentration | 41 | 130/C2/Water | volumetric water concentration
42 | 130/C1/Water | volumetric water concentration | 42 | 130/H/Electric | electric conductivity
43 | 130/C2/Water | volumetric water concentration | 43 | 130/A/Electric | electric conductivity
44 | 130/C3/Water | volumetric water concentration | 44 | 130/B1/Electric | electric conductivity
45 | 130/H/Electric | electric conductivity | 45 | 130/B2/Electric | electric conductivity
46 | 130/A/Electric | electric conductivity | 46 | 130/C1/Electric | electric conductivity
47 | 130/B1/Electric | electric conductivity | 47 | 130/C2/Electric | electric conductivity
48 | 130/B2/Electric | electric conductivity | 48 | 160/H/Water | volumetric water concentration
49 | 130/C1/Electric | electric conductivity | 49 | 160/A/Water | volumetric water concentration
50 | 130/C2/Electric | electric conductivity | 50 | 160/B1/Water | volumetric water concentration
51 | 130/C3/Electric | electric conductivity | 51 | 160/B2/Water | volumetric water concentration
52 | 160/H/Water | volumetric water concentration | 52 | 160/C1/Water | volumetric water concentration
53 | 160/A/Water | volumetric water concentration | 53 | 160/C2/Water | volumetric water concentration
54 | 160/B1/Water | volumetric water concentration | 54 | 160/H/Electric | electric conductivity
55 | 160/B2/Water | volumetric water concentration | 55 | 160/A/Electric | electric conductivity
56 | 160/C1/Water | volumetric water concentration | 56 | 160/B1/Electric | electric conductivity
57 | 160/C2/Water | volumetric water concentration | 57 | 160/B2/Electric | electric conductivity
58 | 160/C3/Water | volumetric water concentration | 58 | 160/C1/Electric | electric conductivity
59 | 160/C4/Water | volumetric water concentration | 59 | 160/C2Electric | electric conductivity
60 | 160/H/Electric | electric conductivity | 60 | 60/A/Water | volumetric water concentration
61 | 160/A/Electric | electric conductivity | 61 | 60/B/Water | volumetric water concentration
62 | 160/B1/Electric | electric conductivity | 62 | 60/C/Water | volumetric water concentration
63 | 160/B2/Electric | electric conductivity | 63 | 60/A/Electric | electric conductivity
64 | 160/C1/Electric | electric conductivity | 64 | 60/B/Electric | electric conductivity
65 | 160/C2Electric | electric conductivity | 65 | 60/C/Electric | electric conductivity
66 | 160/C3/Electric | electric conductivity | 66 | 80/A/Water | volumetric water concentration
67 | 160/C4/Electric | electric conductivity | 67 | 80/B1/Water | volumetric water concentration
68 | WEIR1/32/Water | volumetric water concentration | 68 | 80/B2/Water | volumetric water concentration
69 | WEIR1/33/Water | volumetric water concentration | 69 | 80/C/Water | volumetric water concentration
70 | WEIR1/34/Water | volumetric water concentration | 70 | 80/A/Electric | electric conductivity
71 | WEIR1/35/Water | volumetric water concentration | 71 | 80/B1/Electric | electric conductivity
72 | WEIR1/36/Water | volumetric water concentration | 72 | 80/B2/Electric | electric conductivity
73 | WEIR1/37/Water | volumetric water concentration | 73 | 80/C/Electric | electric conductivity
74 | WEIR1/A/Water | volumetric water concentration
75 | WEIR1/B/Water | volumetric water concentration
76 | WEIR1/32/Electric | electric conductivity
77 | WEIR1/33/Electric | electric conductivity
78 | WEIR1/34/Electric | electric conductivity
79 | WEIR1/35/Electric | electric conductivity
80 | WEIR1/36/Electric | electric conductivity
81 | WEIR1/37/Electric | electric conductivity
82 | WEIR1/A/Electric | electric conductivity
83 | WEIR1/B/Electric | electric conductivity
84 | WEIR2/40/Water | volumetric water concentration
85 | WEIR2/41/Water | volumetric water concentration
86 | WEIR2/42/Water | volumetric water concentration
87 | WEIR2/43/Water | volumetric water concentration
88 | WEIR2/44/Water | volumetric water concentration
89 | 60/A/Water | volumetric water concentration
90 | 60/B/Water | volumetric water concentration
91 | 60/C/Water | volumetric water concentration
92 | WEIR2/40/Electric | electric conductivity
93 | WEIR2/41/Electric | electric conductivity
94 | WEIR2/42/Electric | electric conductivity
95 | WEIR2/43/Electric | electric conductivity
96 | WEIR2/44/Electric | electric conductivity
97 | 60/A/Electric | electric conductivity
98 | 60/B/Electric | electric conductivity
99 | 60/C/Electric | electric conductivity
100 | 80/A/Water | volumetric water concentration
101 | 80/B1/Water | volumetric water concentration
102 | 80/B2/Water | volumetric water concentration
103 | 80/C/Water | volumetric water concentration
104 | 80/A/Electric | electric conductivity

## TDR100
TDR100 13.01.2005 |  |  |  | New TDR 100 was started in January 2005. Old calibrations are not valid any more. |  |  | New TDR sensors were installed and calibration was made during 2011-2012. TDR calibration data can be found from dutyinfo on Timo\smear\TDR-calib.xls
Sampo\D:\smear\tdr100\TDYYMMDD..gmo |  |  |  | Data can be found in folder TDR 100. Values are K-values, which have to be converted
to volumetric water content using appropriate calibration functions.
Column | MUX | Label | Measurement | 13.1.-25.5.2005 | Since 26.5.2005 | Since 18.7.2011 | Since 25.8.2012
1 |  | date and time | yyyymmddhhmmss | date and time | date and time | date and time | date and time
2 | 1 | 50/H/Water | volumetric water concentration | 50/H/Water | 50/H/Water | 50/H/Water | 50/H/Water
3 | 1 | 50/A/Water | volumetric water concentration | 50/A/Water | 50/A/Water | 50/A/Water | 50/A/Water
4 | 1 | 50/B/Water | volumetric water concentration | 50/B/Water | 50/B/Water | 50/B/Water | 50/B/Water
5 | 1 | 50/C/Water | volumetric water concentration | 50/C/Water | 50/C/Water | 50/C/Water | 50/C/Water
6 |  | Empty |  | Empty | Empty | Empty | Empty
7 |  | Empty |  | Empty | Empty | Empty | Empty
8 |  | Empty |  | Empty | Empty | Empty | Empty
9 |  | Empty |  | Empty | Empty | Empty | Empty
10 | 2 | 70/H/Water | volumetric water concentration | 70/H/Water | 70/H/Water | 70/H/Water | 70/H/Water
11 | 2 | 70/A/Water | volumetric water concentration | 70/A/Water | 70/A/Water | 70/A/Water | 70/A/Water
12 | 2 | 70/B1/Water | volumetric water concentration | 70/B1/Water | 70/B1/Water | 70/B1/Water | 70/B1/Water
13 | 2 | 70/B2/Water | volumetric water concentration | 70/B2/Water | 70/B2/Water | 70/B2/Water | 70/B2/Water
14 | 2 | 70/C/Water | volumetric water concentration | 70/C/Water | 70/C/Water | 70/C/Water | 70/C/Water
15 |  | Empty |  | Empty | Empty | Empty | Empty
16 | 2 | humus near cyv 142 | volumetric water concentration | humus near cyv 142 23.5.05-> | humus near cyv 142 23.5.05-> | humus near cyv 142 23.5.05-> | humus near cyv 142 23.5.05->
17 | 2 | humus near cyv 142 | volumetric water concentration | humus near cyv 142 23.05.05-> | humus near cyv 142 23.5.05-> | humus near cyv 142 23.5.05-> | Empty
18 | 3 | 100/H/Water | volumetric water concentration | 100/H/Water | 100/H/Water | 100/H/Water | 100/H/Water
19 | 3 | 100/A/Water | volumetric water concentration | 100/A/Water | 100/A/Water | 100/A/Water | 100/A/Water
20 | 3 | 100/B1/Water | volumetric water concentration | 100/B1/Water | 100/B1/Water | 100/B1/Water | 100/B1/Water
21 | 3 | 100/B2/Water | volumetric water concentration | 100/B2/Water | 100/B2/Water | 100/B2/Water | 100/B2/Water
22 | 3 | 100/C/Water | volumetric water concentration | 100/C/Water | 100/C/Water | 100/C/Water | 100/C/Water
23 | 3 | 100/C2/Water | volumetric water concentration | 100/C2/Water | 100/C2/Water | 100/C2/Water | 100/C2/Water
24 | 3 | 100/C3/Water | volumetric water concentration | 100/C3/Water | 100/C3/Water | 100/C3/Water | 100/C3/Water
25 | 3 | 100/H/Water | volumetric water concentration | Empty | Empty | 100/H/Water | 100/H/Water
26 | 4 | 130/H/Water | volumetric water concentration | 130/H/Water | 130/H/Water | 130/H/Water | 130/H/Water
27 | 4 | 130/A/Water | volumetric water concentration | 130/A/Water | 130/A/Water | 130/A/Water | 130/A/Water
28 | 4 | 130/B1/Water | volumetric water concentration | 130/B1/Water | 130/B1/Water | 130/B1/Water | 130/B1/Water
29 | 4 | 130/B2/Water | volumetric water concentration | 130/B2/Water | 130/B2/Water | 130/B2/Water | 130/B2/Water
30 | 4 | 130/C1/Water | volumetric water concentration | 130/C1/Water | 130/C1/Water | 130/C1/Water | 130/C1/Water
31 | 4 | 130/C2/Water | volumetric water concentration | 130/C2/Water | 130/C2/Water | 130/C2/Water | 130/C2/Water
32 | 4 | 130/C3/Water | volumetric water concentration | 130/C3/Water | 130/C3/Water | 130/C3/Water | 130/C3/Water
33 |  | Empty |  | Empty | Empty | Empty | Empty
34 | 5 | 160/H/Water | volumetric water concentration | 160/H/Water | 160/H/Water | 160/H/Water | 160/H/Water
35 | 5 | 160/A/Water | volumetric water concentration | 160/A/Water | 160/A/Water | 160/A/Water | 160/A/Water
36 | 5 | 160/B1/Water | volumetric water concentration | 160/B1/Water | 160/B1/Water | 160/B1/Water | 160/B1/Water
37 | 5 | 160/B2/Water | volumetric water concentration | 160/B2/Water | 160/B2/Water | 160/B2/Water | 160/B2/Water
38 | 5 | 160/C1/Water | volumetric water concentration | 160/C1/Water | 160/C1/Water | 160/C1/Water | 160/C1/Water
39 | 5 | 160/C2/Water | volumetric water concentration | 160/C2/Water | 160/C2/Water | 160/C2/Water | 160/C2/Water
40 | 5 | 160/C3/Water | volumetric water concentration | 160/C3/Water | 160/C3/Water | 160/C3/Water | 160/C3/Water
41 | 5 | 160/C4/Water | volumetric water concentration | 160/C4/Water | 160/C4/Water | 160/C4/Water | 160/C4/Water
42 | 6 | WEIR1/32/Water | volumetric water concentration | WEIR1/32/Water | WEIR1/32/Water | WEIR1/32/Water | WEIR1/32/Water
43 | 6 | WEIR1/33/Water | volumetric water concentration | WEIR1/33/Water | WEIR1/33/Water | WEIR1/33/Water | WEIR1/33/Water
44 | 6 | WEIR1/34/Water | volumetric water concentration | WEIR1/34/Water | WEIR1/34/Water | WEIR1/34/Water | WEIR1/34/Water
45 | 6 | WEIR1/35/Water | volumetric water concentration | WEIR1/35/Water | WEIR1/35/Water | WEIR1/35/Water | WEIR1/35/Water
46 | 6 | WEIR1/36/Water | volumetric water concentration | WEIR1/36/Water | WEIR1/36/Water | WEIR1/36/Water | WEIR1/36/Water
47 | 6 | WEIR1/37/Water | volumetric water concentration | WEIR1/37/Water | WEIR1/37/Water | WEIR1/37/Water | WEIR1/37/Water
48 | 6 | WEIR1/A/Water | volumetric water concentration | WEIR1/A/Water | WEIR1/A/Water | WEIR1/A/Water | WEIR1/A/Water
49 | 6 | WEIR1/B/Water | volumetric water concentration | WEIR1/B/Water | WEIR1/B/Water | WEIR1/B/Water | WEIR1/B/Water
50 | 7 | WEIR2/40/Water | volumetric water concentration | WEIR2/40/Water | WEIR2/40/Water | WEIR2/40/Water | WEIR2/40/Water
51 | 7 | WEIR2/41/Water | volumetric water concentration | WEIR2/41/Water | WEIR2/41/Water | WEIR2/41/Water | WEIR2/41/Water
52 | 7 | WEIR2/42/Water | volumetric water concentration | WEIR2/42/Water | WEIR2/42/Water | WEIR2/42/Water | WEIR2/42/Water
53 | 7 | WEIR2/43/Water | volumetric water concentration | WEIR2/43/Water | WEIR2/43/Water | WEIR2/43/Water | WEIR2/43/Water
54 | 7 | WEIR2/44/Water | volumetric water concentration | WEIR2/44/Water | 60 Humus/Water | 61 Humus/Water | 62 Humus/Water
55 | 7 | 60/A/Water | volumetric water concentration | 60/A/Water | 60/A/Water | 60/A/Water | 60/A/Water
56 | 7 | 60/B/Water | volumetric water concentration | 60/B/Water | 60/B/Water | 60/B/Water | 60/B/Water
57 | 7 | 60/C/Water | volumetric water concentration | 60/C/Water | 60/C/Water | 60/C/Water | 60/C/Water
58 | 8 | 80/A/Water | volumetric water concentration | 80/A/Water | 80/A/Water | 80/A/Water | 80/A/Water
59 | 8 | 80/B1/Water | volumetric water concentration | 80/B1/Water | 80/B1/Water | 80/B1/Water | 80/B1/Water
60 | 8 | 80/B2/Water | volumetric water concentration | 80/B2/Water | 80/B2/Water | 80/B2/Water | 80/B2/Water
61 | 8 | 80/C/Water | volumetric water concentration | 80/C/Water | 80/C/Water | 80/C/Water | 80/C/Water
62 |  | Empty |  | Empty | Empty | Empty | Empty
63 |  | Empty |  | Empty | Empty | Empty | Empty
64 | 8 | Humus | volumetric water concentration | Humus | 80/Humus/Water | 80/Humus/Water | Empty
65 | 8 | Humus | volumetric water concentration | Humus | 80/Humus/Water | 80/Humus/Water | Empty

## Tdmon
Louhi\D:\smear\Tdmon\SMYYMMDD.tdmon
Began 1.12.2011
Datafile column | Nokeval 13 channel | label | units
1 |  | date and time | yyyymmddhhmmss
2 | 5 | Dew point temperature | °C
3 | 6 | Ambient temperature | °C
4 | 7 | Flow status
Nokeval 11 channel
5 | 1 | Ambient pressure | hPa
Began 15.6.2012
Datafile column | Nokeval 13 channel | label | units
1 |  | date and time | yyyymmddhhmmss
2 | 5 | Dew point temperature | °C
3 | 6 | Ambient temperature | °C
4 | 7 | Flow status | mV
5 | 9 | Ambient RH (Rotronic) | mV
6 | 10 | Ambient temperature (Rotronic) | mV
7 | 11 | Ambient temperature (PT100) | °C
Nokeval 11 channel
8 | 1 | Ambient pressure | hPa

## Eddymon
Louhi\D:\smear\eddymon\SMYYMMDD.EMON
Began 15.03.2002
Datafile column | Nokeval 13 channel | label | units
1 |  | date and time | yyyymmddhhmmss
2 | 1 | LOZ-3 internal pressure | hPa
3 | 2 | LOZ-3 chamber temperature | °C
4 | 3 | LOZ-3 liquid flow | ml/min
Began 14.05.2010
Datafile column | Nokeval 13 channel | label | units
1 |  | date and time | yyyymmddhhmmss
2 | 1 | LOZ-3 internal pressure | hPa
3 | 2 | LOZ-3 chamber temperature | °C
4 | 3 | Pressure difference, dP | hPa

## Eddymet
SMEAR II | Column order in Eddy-datafiles (Smyymmdd_*.flx)
Pepe\D:\smetek\sYYMMDD_MET.flx
Valid 05.02.2004 - (see Notes for some details)
Datafile column | Measured Quantity | Unit in datafile | Description | Notes
1 | Yr | yy | measurement year
2 | Day | day number | number of day from year's beginning
3 | h | hh | starting hour of 1/2 an hour period
4 | m | m or mm | ending minute of 1/2 an hour period
5 | Sonic_H | W·m-2 | sensible heat flux | neg. sign downwards
6 | LE | W·m-2 | latent heat flux | neg. sign downwards
7 | E | mmol·m-2 s-1 | water vapour flux | neg. sign downwards
8 | F_c | µmol·m-2 s-1 | carbon dioxide flux | neg. sign downwards
9 | tau | m2·s-2 | momentum flux | neg. sign downwards
10 | u* | m·s-1 | friction velocity
11 | MO_length | m | Monin-Obukhov length
12 | std_u | m·s-1 | st. deviation of wind component u
13 | std_v | m·s-1 | st. deviation of wind component v
14 | std_w | m·s-1 | st. deviation of wind component w
15 | std_t | °C | st. deviation of air temperature
16 | std_c | ppm | st. deviation of CO2 concentration
17 | std_h | mmol·mol-1 | st. deviation of H2O concentration
18 | av u | m·s-1 | average wind component u
19 | av v | m·s-1 | average wind component v
20 | av w | m·s-1 | average wind component w
21 | av t | °C | average air temperature
22 | av_c | ppm | average CO2 concentration
23 | av_h | mmol·mol-1 | average H2O concentration
24 | av U | m·s-1 | average horiz. wind (u & v combined)
25 | wind_dir | ° | average wind direction
26 | Eta | ° | co-ordinate rotation angle
27 | Theta | ° | co-ordinate rotation angle
28 | Beta | ° | co-ordinate rotation angle
29 | c_lag | 1/20 s | CO2 concentration lag time
30 | h_lag | 1/20 s | H2O concentration lag time
31 | err_packet | number | number of erroneous data packets
32 | cov_wAn3
33 | std_An3
34 | av_An3
35 | lag_An3
36 | cov_wAn4
37 | std_An4
38 | av_An4
39 | lag_An4
40 | cov_wAn5
41 | std_An5
42 | av_An5
43 | lag_An5
44 | cov_wAn6
45 | std_An6
46 | av_An6
47 | lag_An6
48 | cov_wAn7
49 | std_An7
50 | av_An7
51 | lag_An7
52 | cov_wAn8
53 | std_An8
54 | av_An8
55 | lag_An8

## Eddytow
SMEAR II |  |  | Column order in tower Eddy-datafiles (Smyymmdd_tow.flx)
Eddy\D:\eddytow_data\fluxdata\sYYMMDD_tow.flx
Valid 01.08.2001 - (see Notes for some details)
Datafile column | Measured Quantity | Unit in datafile | Description | Notes
1 | Yr | yy | measurement year
2 | Day | day number | number of day from year’s beginning
3 | h | hh | starting hour of ½ an hour period
4 | m | m or mm | ending minute of ½ an hour period
5 | Sonic_H | W·m-2 | sensible heat flux | neg. sign downwards
6 | LE | W·m-2 | latent heat flux | neg. sign downwards
7 | E | mmol·m-2 s-1 | water vapour flux | neg. sign downwards
8 | F_c | mol·m-2 s-1 | carbon dioxide flux | neg. sign downwards
9 | tau | m2·s-2 | momentum flux | neg. sign downwards
10 | u* | m·s-1 | friction velocity
11 | MO_length | m | Monin-Obukhov length
12 | std_u | m·s-1 | st. deviation of wind component u
13 | std_v | m·s-1 | st. deviation of wind component v
14 | std_w | m·s-1 | st. deviation of wind component w
15 | std_t | C | st. deviation of air temperature
16 | std_c | ppm | st. deviation of CO2 concentration
17 | std_h | mmol·mol-1 | st. deviation of H2O concentration
18 | av u | m·s-1 | average wind component u
19 | av v | m·s-1 | average wind component v
20 | av w | m·s-1 | average wind component w
21 | av t | C | average air temperature
22 | av_c | ppm | average CO2 concentration
23 | av_h | mmol·mol-1 | average H2O concentration
24 | av U | m·s-1 | average horiz. wind (u & v combined)
25 | wind_dir |  | average wind direction
26 | Eta |  | co-ordinate rotation angle
27 | Theta |  | co-ordinate rotation angle
28 | Beta |  | co-ordinate rotation angle
29 | c_lag | 1/10 s | CO2 concentration lag time
30 | h_lag | 1/10 s | H2O concentration lag time
31 | err_packet | number | number of erraneous data packets
32 | covar_T_L-w | C·m·s-1 | non data
33 | std_T_L | C | st. deviation of LiCor temperature
34 | av_T_L | C | average LiCor temperature
35 | T_L_lag | 1/10 s | non data
36 | covar_P_L-w | kPa·m·s-1 | non data
37 | std_p_L | kPa | st. deviation of LiCor sample pressure
38 | av_p_L | kPa | average LiCor sample pressure
39 | p_L_lag | 1/10 s | non data
40 | F_O3 | ppb·m·s-1 | O3 flux | neg.sign downwards
41 | std_O3 | ppb | st. deviation of O3 concentration
42 | av_O3 | ppb | average O3 concentration | O3 analys. connected 06.08.2001
43 | O3_lag | 1/10 s | O3 concentration lag time
44 | covar_LOZ_F-w | l·min-1·m·s-1 | non data
45 | std_LOZ_F | l·min-1 | st.deviation of LOZ sample flow
46 | av_LOZ_F | l·min-1 | average LOZ sample flow
47 | LOZ_F_lag | 1/10 s | non data

## REA
SMEAR II |  |  | Column order in REA-datafiles (syymmdd_REA.dat)
REA\C:\rea\sYYMMDD_REA.dat
Datafile column | Measured Quantity | Unit in datafile | Description | Notes
1 | Yr | yy | measurement year
2 | Day | day number | number of day from years beginning
3 | h | hh | starting hour of 1/2 an hour period
4 | m | m or mm | ending minute of 1/2 an hour period
5 | Sonic_H | W·m-2 | sensible heat flux | neg. sign downwards
6 | uw |  | uw
7 | u* | m·s-1 | friction velocity
8 | MO_length | m | Monin-Obukhov length
9 | std_u | m·s-1 | st. deviation of wind component u
10 | std_v | m·s-1 | st. deviation of wind component v
11 | std_w | m·s-1 | st. deviation of wind component w
12 | std_t | C | st. deviation of air temperature
13 | av_u | m·s-1 | average wind component u
14 | av_v | m·s-1 | average wind component v
15 | av_w | m·s-1 | average wind component w
16 | av t | C | average air temperature
17 | av_U | m·s-1 | average horiz. wind (u & v combined)
18 | wind_dir |  | average wind direction
19 | Eta |  | co-ordinate rotation angle
20 | Theta |  | co-ordinate rotation angle
21 | Beta |  | co-ordinate rotation angle
22 | err_packet | number | number of erraneous data packets
23 | %Up |  | %Up
24 | %Down |  | %Down
25 | %Dead |  | %Dead
26 | Tup |  | Tup
27 | Tdown |  | Tdown
28 | Tdead |  | Tdead
29 | dTup |  | dTup
30 | dTdown |  | dTdown
31 | dTdead |  | dTdead
32 | H_REA |  | H_REA
33 | Uup |  | Uup
34 | Udown |  | Udown
35 | Udead |  | Udead
36 | dUup |  | dUup
37 | dUdown |  | dUdown
38 | dUdead |  | dUdead
39 | wu_REA |  | wu_REA
40 | Vup |  | Vup
41 | Vdown |  | Vdown
42 | Vdead |  | Vdead
43 | dVup |  | dVup
44 | dVdown |  | dVdown
45 | dVdead |  | dVdead
46 | wv_REA |  | wv_REA

## Eddy233
SMEAR II | Column order in Eddy-datafiles (Smyymmdd_*.flx)
Eddy\D:\eddy233_data\fluxdata\sYYMMDD_233.flx
Valid 01.08.2000 - (see Notes for some details)
Datafile column | Measured Quantity | Unit in datafile | Description | Notes
1 | Yr | yy | measurement year
2 | Day | day number | number of day from year's beginning
3 | h | hh | starting hour of 1/2 an hour period
4 | m | m or mm | ending minute of 1/2 an hour period
5 | Sonic_H | W·m-2 | sensible heat flux | neg. sign downwards
6 | LE | W·m-2 | latent heat flux | neg. sign downwards
7 | E | mmol·m-2 s-1 | water vapour flux | neg. sign downwards
8 | F_c | µmol·m-2 s-1 | carbon dioxide flux | neg. sign downwards
9 | tau | m2·s-2 | momentum flux | neg. sign downwards
10 | u* | m·s-1 | friction velocity
11 | MO_length | m | Monin-Obukhov length
12 | std_u | m·s-1 | st. deviation of wind component u
13 | std_v | m·s-1 | st. deviation of wind component v
14 | std_w | m·s-1 | st. deviation of wind component w
15 | std_t | °C | st. deviation of air temperature
16 | std_c | ppm | st. deviation of CO2 concentration
17 | std_h | mmol·mol-1 | st. deviation of H2O concentration
18 | av u | m·s-1 | average wind component u
19 | av v | m·s-1 | average wind component v
20 | av w | m·s-1 | average wind component w
21 | av t | °C | average air temperature
22 | av_c | ppm | average CO2 concentration
23 | av_h | mmol·mol-1 | average H2O concentration
24 | av U | m·s-1 | average horiz. wind (u & v combined)
25 | wind_dir | ° | average wind direction
26 | Eta | ° | co-ordinate rotation angle
27 | Theta | ° | co-ordinate rotation angle
28 | Beta | ° | co-ordinate rotation angle
29 | c_lag | 1/20 s | CO2 concentration lag time
30 | h_lag | 1/20 s | H2O concentration lag time
31 | err_packet | number | number of erroneous data packets
32 | covar_T_L-w | °C·m·s-1 | non data
33 | std_T_L | °C | st. deviation of LiCor temperature
34 | av_T_L | °C | average LiCor temperature
35 | T_L_lag | 1/20 s | non data
36 | covar_P_L-w | kPa·m·s-1 | non data
37 | std_p_L | kPa | st. deviation of LiCor sample pressure
38 | av_p_L | kPa | average LiCor sample pressure
39 | p_L_lag | 1/20 s | non data
40 | F_CPC | 109 m-2·s-1 | particle flux | neg.sign downwards
41 | std_CPC | 103 m-3 | st. deviation of particle concentration
42 | av_CPC | 103 m-3 | average particle concentration | CPC connected 01.08.2000
43 | CPC_lag | 1/20 s | CPC concentration lag time
44 | empty |  | reserved
45 | empty |  | reserved
46 | empty |  | reserved
47 | empty |  | reserved

## Eddy460
File syymmdd_460.flx
Column | Meaning | Unit
1 | Year
2 | Day | julian day
3 | Hour
4 | Minute
5 | Sonic_H | W/m2
6 | LE | W/m2
7 | E | mmol/(m2s)
8 | F_c | µmol/(m2s)
9 | tau | m2/s2
10 | u* | m/s
11 | MO_length | m
12 | std_u | m2/s2
13 | std_v | m2/s2
14 | std_w | m2/s2
15 | std_t | (°C)2
16 | std_c | ppm2
17 | std_h | mmol/mol
18 | av_u | m/s
19 | av_v | m/s
20 | av_w | m/s
21 | av_t | °C
22 | av_c | ppm
23 | av_h | ppth
24 | U | m/s
25 | wind_dir | degrees
26 | Eta | degrees
27 | Theta | degrees
28 | Beta | degrees
29 | c_lag | sample periods
30 | h_lag | sample periods
31 | err_packet | number
32 | covar_T_L-w
33 | covar_P_L-w
34 | F_CPC | 109 particles/(m2s)
35 | var_T_L | (°C)2
36 | var_P_L | (kPa)2
37 | var_CPC | (103 particles/cm3)2
38 | av_T_L | °C
39 | av_P_L | kPa
40 | av_CPC | 103 particles/cm3
41 | T_L lag
42 | P_L lag
43 | CPC_lag | sample periods

## Heat
Louhi\D:\smear\heat\SMYYMMDD.HEA |  |  |  | Began 23.11.1999
Column | Pepe Com 6, Iiro Com 5 | ProMux Channel | Meaning | Unit in datafile
1 |  |  | date and time | yyyymmddhhmmss
2 |  |  | moment | s
3 |  | 0 | 1 A | °C
4 |  | 1 | 1 B | °C
5 |  | 2 | 2 A | °C
6 |  | 3 | 2 B | °C
7 |  | 4 | 3 A | °C
8 |  | 5 | 3 B | °C

## Lautta
SMEAR II, Lautta |  |  | Valid 8.06.2012 =>4.4.2017
Heikki2\D:\progs\nokng\data\Smyymmdd.LA1
Datafile column | Nokeval Channel | Measurement | Units of measurement | Descriptions
1 |  | date& time | yyyymmddhhmmss | depth
Nokeval 0
2 | 1 | Temperature | °C | 0,2m
3 | 2 | Temperature | °C | 0,5m
4 | 3 | Temperature | °C | 1,0m
5 | 4 | Temperature | °C | 1,5m
6 | 5 | Temperature | °C | 2,0m
7 | 6 | Temperature | °C | 2,5m
8 | 7 | Temperature | °C | 3,0m
9 | 8 | Temperature | °C | 3,5m
10 | 9 | Temperature | °C | 4,0m
11 | 10 | Temperature | °C | 4,5m
12 | 11 | Temperature | °C | 5,0m
13 | 12 | Temperature | °C | 6,0m
14 | 13 | Temperature | °C | 7,0m
15 | 14 | Temperature | °C | 8,0m
16 | 15 | Temperature | °C | 10,0m
17 | 16 | Temperature | °C | 12,0m
Nokeval 1
18 | 1 | CM 3 up | mV | Radiation
19 | 2 | CM 3 down | mV | Radiation
20 | 3 | CG 3 up | mV | Radiation
21 | 4 | CG 3 down | mV | Radiation
22 | 5 | Pt100 | mV | Temperature
23 | 6
24 | 7 | T amb. (Rotronic) | mV
25 | 8 | RH amb. (Rotronic) | mV
26 | 9 | Vaisala CO2 |  | 0.2m
27 | 10 | Vaisala CO2 |  | 2.5m
28 | 11 | Vaisala CO2 |  | 1.5m
29 | 12 | Pt100 | °C | Temperature
30 | 13 | Li-193 | mV | Underwater radiation (0.2 m)
31 | 14 | Li-192 | mV | Underwater radiation (0.5 m)
32 | 15 | Li-192 | mV | Underwater radiation (1.0 m)
33 | 16 | Li-192 | mV | Underwater radiation (2.0 m)

## Mastrad
SMEAR II |  | Column order in mastrad-datafiles (*.mrd)
Valid since 11.11.2009
Sampo\D:\smear\mastrad\SMYYMMDD.mrd
Datafile column | Nokeval 0, Port 21 in Sampo PC | Measured Quantity | Unit in datafile | Description | Notes
1 | channel | date & time | yyyymmddhhmmss
2 | 1 | Skye IR in | mV
3 | 2 | IR ref temp. | °C
4 | 3 | Skye IR out | mV
5 | 4 | IR ref temp. | °C
6 | 5 | NDVI in 1 | mV | up facing | 1 kohm
7 | 6 | NDVI in 2 | mV | up facing | 1 kohm
8 | 7 | NDVI out 1 | mV | down facing, 60 deg collar and diffusor | 10 kohm
9 | 8 | NDVI out 2 | mV | down facing, 60 deg collar and diffusor | 10 kohm
10 | 9 | PRI in 1 | mV | up facing, diffusor | 5 kohm
11 | 10 | PRI in 2 | mV | up facing, diffusor | 5 kohm
12 | 11 | PRI out 1 | mV | down facing | 5 kohm
13 | 12 | PRI out 2 | mV | down facing | 5 kohm
14 | 13 | Empty
15 | 14 | Empty
16 | 15 | Empty
17 | 16 | T box | °C
Valid since 16.12.2009
Sampo\D:\smear\mastrad\SMYYMMDD.mrd
Datafile column | Nokeval 0, Port 21 in Sampo PC | Measured Quantity | Unit in datafile | Description | Notes
1 | channel | date & time | yyyymmddhhmmss
2 | 1 | Skye IR in | mV
3 | 2 | IR ref temp. | °C
4 | 3 | Skye IR out | mV
5 | 4 | IR ref temp. | °C
6 | 5 | NDVI in 1 | mV | up facing | 1 kohm
7 | 6 | NDVI in 2 | mV | up facing | 1 kohm
8 | 7 | NDVI out 1 | mV | down facing, 60 deg collar and diffusor | 10 kohm
9 | 8 | NDVI out 2 | mV | down facing, 60 deg collar and diffusor | 10 kohm
10 | 9 | PRI in 1 | mV | up facing, diffusor | 5 kohm
11 | 10 | PRI in 2 | mV | up facing, diffusor | 5 kohm
12 | 11 | PRI out 1 | mV | down facing | 5 kohm
13 | 12 | PRI out 2 | mV | down facing | 5 kohm
14 | 16 | T box | °C

## MastMet33
SMEAR II, MastMet33 |  |  | Valid 18.01.2013 =>
Gaslog\C:\MastMET\33m_data\SMyymmdd.33m
Datafile column | Nokeval Channel | Measurement | Units of measurement | Descriptions | Notes
1 |  | date& time | yyyymmddhhmmss
Nokeval 33
2 | 1 | IR in | [mV]
3 | 2 | In ref T | [mV]
4 | 3 | IR out | [mV]
5 | 4 | Out ref T | °C
6 | 5 | NDVI in1 652nm |  | up facing | 1 kohm
7 | 6 | NDVI in2 860nm |  | up facing | 1 kohm
8 | 7 | NDVI out1 652nm |  | down facing, no collar, no diffusor | 10 kohm
9 | 8 | NDVI out2 860nm |  | down facing, no collar, no diffusor | 10 kohm
10 | 9 | PRI in1 531nm |  | up facing, diffusor | 5 kohm
11 | 10 | PRI in2 570nm |  | up facing, diffusor | 5 kohm
12 | 11 | PRI out1 531nm |  | down facing | 5 kohm
13 | 12 | PRI out2 570nm |  | down facing | 5 kohm
14 | 13 | T Pt 16m | °C
15 | 14 | T mV 16m | [mV]
16 | 15 | RH 16m | [%]
17 | 16 | Instrument box temperature (Pt100) | [mV]

## MastMet67
SMEAR II, MastMet67 |  |  | Valid 27.08.2012 =>
Gaslog\C:\MastMET\67m_data\SMyymmdd.67m
Datafile column | Nokeval Channel | Measurement | Units of measurement | Descriptions
1 |  | date& time | yyyymmddhhmmss
Nokeval 7
2 | 1 | Temperature Rotronic external (Pt100) | [mV]
3 | 2 | Temperature Rotronic | [mV]
4 | 3 | RH Rotronic | [mV]
5 | 4 | Empty
6 | 5 | Empty
7 | 6 | Empty
8 | 7 | Empty
9 | 8 | Empty
10 | 9 | Empty
11 | 10 | Empty
12 | 11 | Empty
13 | 12 | Empty
14 | 13 | Empty
15 | 14 | Empty
16 | 15 | Empty
17 | 16 | Instrument box temperature (Pt100) | [mV]

## MastMet125
SMEAR II, MastMet125 |  |  | Valid 28.08.2012 =>
Gaslog\C:\MastMET\125m_data\SMyymmdd.67m
Datafile column | Nokeval Channel | Measurement | Units of measurement | Descriptions
1 |  | date& time | yyyymmddhhmmss
Nokeval 125
2 | 1 | Temperature Rotronic external (Pt100) | [mV]
3 | 2 | Temperature Rotronic | [mV]
4 | 3 | RH Rotronic | [mV]
5 | 4 | Up looking pyranometer | [mV]
6 | 5 | Down looking pyranometer | [mV]
7 | 6 | Empty
8 | 7 | Empty
9 | 8 | Empty
10 | 9 | Empty
11 | 10 | Empty
12 | 11 | Empty
13 | 12 | Empty
14 | 13 | Empty
15 | 14 | Empty
16 | 15 | Empty
17 | 16 | Instrument box temperature (Pt100) | [mV]

## Mast 01.03.2007
SMEAR II, MAST |  |  | Valid 26.03.2004 =>
Louhi\D:\smear\mast\SMYYMMDD.MAS
Nokeval, Promux + AD-converter configuration
Datafile column | Nokeval Channel | Measurement | Units of measurement | Descriptions | Notes
1 |  | date& time | Yyyymmddhhmmss
2 |  | height | dm
| Nokeval 15 | port 5 in Louhi PC
3 | 1 | T Enc | °C | Nokeval & Promux enclosure temperature
4 | 2 | Td | °C | Dew point temperature
5 | 3 | Netref | mV | “Reference” net radiation | Sensor disconnected 25.08.03
6 | 4 | SkyIR | °C | Sky temperature | Sensor unconnected
7 | 5 | CanIR | °C | Canopy temperature | Sensor unconnected
8 | 6 | R PAR | mV | Reflected Par
9 | 7 | R K | mV | Reflected Global | sensor disconnected 21.09.2011
10 | 8 | Net | mV | Net Radiation | sensor disconnected 21.09.2011
11 | 9 | Empty
12 | 10 | Empty
13 | 11 | T Eddy CPC box | °C | Eddy CPC box temperature
14 | 12 | Empty
15 | 13 | MFC LiCor Sample | mV | Eddy 233 LiCor sample flow
16 | 14 | Empty
17 | 15 | Empty
18 | 16 | WSU 8.4 N | m/s | Wind speed, ultrasonic, 8.4m, North | North side
| Nokeval 14 | port 5 in Louhi PC
19 | 1 | T 67.2 | °C | Temperature 67.2m
20 | 2 | WDU 74.0 | ° | Wind direction, ultrasonic, 74m | At 67.2m since 11.07.2011
21 | 3 | T 50.4 | °C | Temperature 50.4m
22 | 4 | WDU 33.6 N | ° | Wind direction, ultrasonic, 33.6m, North | North side
23 | 5 | T 33.6 | °C | Temperature 33.6m
24 | 6 | WDU 33.6 S | ° | Wind direction, ultrasonic, 33.6m, South | South side
25 | 7 | T 16.8 | °C | Temperature 16.8m
26 | 8 | WDU 16.8 N | ° | Wind direction, ultrasonic, 16.8m, North | North side
27 | 9 | T 8.4 | °C | Temperature 8.4m
28 | 10 | WDU 8.4 N | ° | Wind direction, ultrasonic, 8.4m, North | North side
29 | 11 | T 4.2 | °C | Temperature 4.2m
30 | 12 | WSU 74.0 | m/s | Wind speed, ultrasonic, 74m | At 67.2m since 11.07.2011
31 | 13 | WSU 33.6 N | m/s | Wind speed, ultrasonic, 33.6m, North | North side
32 | 14 | WSU 33.6 S | m/s | Wind speed, ultrasonic, 33.6m, South | South side
33 | 15 | WSU 16.8 N | m/s | Wind speed, ultrasonic, 16.8m, North | North side
34 | 16 | Ref T0 | °C | PT-100 reference | Nokeval box
| Nokeval 11 | port 8 in Louhi PC
35 | 1 | BAR | hPa | Ambient pressure
36 | 2 | NO Mast | ppb | NO concentration
37 | 3 | NOX Mast | ppb | NOx concentration
38 | 4 | O3 Mast | ppb | O3 concentration
39 | 5 | NO Photo ch. press | hPa
40 | 6 | NOx Photo ch. press | hPa
41 | 7 | SO2 Mast ch.press. | hPa
42 | 8 | NOx Mast ch press | hPa
43 | 9 | CO mast | ppb | CO concentration
44 | 10 | used by Photo
45 | 11 | used by Photo
46 | 12 | used by Photo
47 | 13 | SO2 Mast | ppb | SO2 concentration
48 | 14 | used by Photo
49 | 15 | used by Photo
50 | 16 | O3/Nox/SO2 press | hPa
| ProMux 0 | port 7 in Louhi PC
51 | 0 | Empty
52 | 1 | Empty
53 | 2 | WS 33.6 N | m/s | Wind speed 33.6m, North | North side, cup
54 | 3 | WS 33.6 S | m/s | Wind speed 33.6m, South | South side, cup
55 | 4 | Empty
56 | 5 | Empty
57 | 6 | Empty
58 | 7 | Empty
59 | 8 | Empty
60 | 9 | Empty
61 | 10 | Empty
62 |  | Empty
63 |  | Empty
64 |  | Empty
65 |  | Empty
66 |  | Empty
AD-converter
67 | 0 | used by Photo |  |  | Photo CO2
68 | 1 | used by Photo |  |  | Photo H2O
69 | 2 | H2O Mast | g/m3 | H2O concentration
70 | 3 | CO2 Mast | ppm | CO2 concentration
71 | 4 | URAS Press | hPa
72 | 5 | used by Photo |  |  | Photo CO2+H2O sample press. drop
73 | 6 | used by Photo |  |  | cuvette supply air MFC
74 | 7 | Empty

## Mast 31.10.2003
SMEAR II MAST |  |  | Valid 22.5.1998 - 25.03.2004
Nokeval, Promux + AD-converter configuration
Datafile column | Nokeval Channel | Measurement | Units of measurement | Descriptions | Notes
1 |  | date& time | yyyymmddhhmmss
2 |  | height | Dm
| Nokeval 15 | port 5 in Pepe PC
3 | 1 | T Enc | °C | Nokeval & Promux enclosure temperature
4 | 2 | Td | °C | dew point
5 | 3 | Netref | mV | "Reference" net radiation | Sensor disconnected 25.08.03
6 | 4 | SkyIR | °C | Sky temperature | Sensor unconnected
7 | 5 | CanIR | °C | Canopy temperature | Sensor unconnected
8 | 6 | R PAR | mV | Reflected Par
9 | 7 | R K | mV | Reflected Global
10 | 8 | Net | mV | Net Radiation
11 | 9 | MFC_S 460 |  | Eddy 460 Licor sample flow | Until 28.6.2000
12 | 10 | Empty
13 | 11 | T Eddy CPC | °C | Eddy CPC box temperature | Began 1.8.2000
14 | 12 | Empty
15 | 13 | MFC_S 233 | mV | Eddy 233 Licor sample flow
16 | 14 | MFC_R |  | Eddy Licor reference flow
17 | 15 | P_L | mV
18 | 16 | T_L | mV
| Nokeval 14 | port 5 in Pepe PC
19 | 1 | T 67.2 | °C
20 | 2 | Empty
21 | 3 | T 50.4 | °C
22 | 4 | WD 50.4 | °
23 | 5 | T 33.6 | °C
24 | 6 | WD 33.6 | °
25 | 7 | T 16.8 | °C
26 | 8 | WD 16.8 | °
27 | 9 | T 8.4 | °C
28 | 10 | Empty
29 | 11 | T 4.2 I A | °C |  | Sensor unconnected
30 | 12 | T 4.2 II B | °C
31 | 13 | Empty
32 | 14 | Empty
33 | 15 | Empty
34 | 16 | Empty
| Nokeval 11 | port 8 in Pepe PC
35 | 1 | BAR | hPa
36 | 2 | NO Mast | ppb
37 | 3 | NOX Mast | ppb
38 | 4 | O3 Mast | ppb
39 | 5 | used by Photo
40 | 6 | used by Photo
41 | 7 | used by Photo
42 | 8 | MFC_URAS | ml/min
43 | 9 | Empty
44 | 10 | NOx Photo | ppb
45 | 11 | URAS Press | hPa
46 | 12 | NO Photo | ppb
47 | 13 | SO2 Mast | ppb
48 | 14 | O3 Photo | ppb
49 | 15 | used by Photo
50 | 16 | O3/Nox/SO2 press | hPa
| ProMux 0 | port 7 in Pepe PC
51 | 0 | WS 67.2 | m/s | North side
52 | 1 | WS 50.4 | m/s | North side
53 | 2 | WS 33.6 A | m/s | North side
54 | 3 | WS 33.6 B | m/s | South side
55 | 4 | Empty
56 | 5 | Empty
57 | 6 | Empty
58 | 7 | Empty
59 | 8 | WS 16.8 | m/s | North side | Until 3.9.2003
60 | 9 | WS 8.4 | m/s | North side | Until 3.9.2003
61 | 10 | WS 4.2 | m/s | North side | Until 3.9.2003
62 |  | Empty
63 |  | Empty
64 |  | Empty
65 |  | Empty
66 |  | Empty
AD-converter
67 | 0 | used by Photo
68 | 1 | used by Photo
69 | 2 | H2O Mast | g/m3 | H2O concentration | Began 29.6.1998
70 | 3 | CO2 Mast | ppm | CO2 concentration | Began 29.6.1998
71 | 4 | URAS Press | hPa |  | Began 29.6.1998
72 | 5 | used by Photo
73 | 6 | used by Photo
74 | 7 | Empty

## Mast 03.12.1995
SMEAR II MAST |  |  | Valid 03.12.1995 - 21.05.1998
Nokeval, Promux + AD-converter configuration
Datafile column | Nokeval Channel | Measurement | Units of measurement | Descriptions | Notes
1 |  | date& time | yyyymmddhhmmss
2 |  | height | Dm
| Nokeval 15 | port 5 in Pepe PC
3 | 1 | T Enc | °C | Nokeval & Promux enclosure temperature
4 | 2 | Empty
5 | 3 | Empty
6 | 4 | Empty
7 | 5 | Empty
8 | 6 | R PAR | mV | Reflected Par
9 | 7 | R K | mV | Reflected Global
10 | 8 | Net | mV | Net Radiation
11 | 9 | MFC_S 460 |  | Eddy 460 Licor sample flow | Began 20.2.1998
12 | 10 | Empty
13 | 11 | Empty
14 | 12 | Empty
15 | 13 | MFC_S 233 | mV | Eddy 233 Licor sample flow
16 | 14 | MFC_R |  | Eddy Licor reference flow
17 | 15 | P_L | mV |  | Began 9.1.1998
18 | 16 | T_L | mV |  | Began 9.1.1998
| Nokeval 14 | port 5 in Pepe PC
19 | 1 | T 67.2 | °C
20 | 2 | Empty
21 | 3 | T 50.4 | °C
22 | 4 | WD 50.4 | °
23 | 5 | T 33.6 | °C
24 | 6 | WD 33.6 | °
25 | 7 | T 16.8 | °C
26 | 8 | WD 16.8 | °
27 | 9 | T 8.4 | °C
28 | 10 | Empty
29 | 11 | Empty |  |  | Sensor unconnected
30 | 12 | T 4.2 II B | °C
31 | 13 | Empty
32 | 14 | Empty
33 | 15 | Empty
34 | 16 | Empty
| Nokeval 11 | port 8 in Pepe PC
35 | 1 | BAR | hPa
36 | 2 | NO Mast | ppb
37 | 3 | NOX Mast | ppb
38 | 4 | O3 Mast | ppb |  | Began 1.8.1996
39 | 5 | used by Photo
40 | 6 | used by Photo
41 | 7 | used by Photo
42 | 8 | MFC_URAS | ml/min
43 | 9 | H2S | ppb |  | Between 28.3.1997 - 14.11.1997
44 | 10 | Empty
45 | 11 | URAS Press | hPa
46 | 12 | Empty
47 | 13 | SO2 Mast | ppb
48 | 14 | Empty
49 | 15 | H2O | g/m3
50 | 16 | CO2 | ppm
| ProMux 0 | port 7 in Pepe PC
51 | 0 | WS 67.2 | m/s | North side
52 | 1 | WS 50.4 | m/s | North side
53 | 2 | WS 33.6 A | m/s | North side
54 | 3 | WS 33.6 B | m/s | South side
55 | 4 | Empty
56 | 5 | Empty
57 | 6 | Empty
58 | 7 | Empty
59 | 8 | WS 16.8 | m/s | North side
60 | 9 | WS 8.4 | m/s | North side
61 | 10 | WS 4.2 | m/s | North side
62 |  | Empty
63 |  | Empty
64 |  | Empty
65 |  | Empty
66 |  | Empty

## NO-monitor_torni
LOGGING DATA FROM Eco Physics (ref. to as ecop hereafter) CLD 780 TR NO-ANALYSER
Program 'ecop_log2.exe' in d:\towergas is used to log the data.
Start the logging by double-clicking the (short-cut) batch file icon 'NOmonitor'
on the desktop.
- it reads ini-file 'ecop.ecm' in folder ...\towergas\data
- data is written to file 'epYYYYMMDD.dat' in folder ...\towergas\data
To stop the logging and exit the program press Ctrl-C in the command window, and
confirm.
ini-file 'ecop.ecm' description:
- # means comment line; works only if the first character of a line; don't insert
any comments after the commands
- don't add extra empty lines at the end of the file
%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%
ini-file 'ecop.ecm'
line: description:
# ecop
COMPORT com3 serial port number
#COMPORT \\.\com3
SEPARATOR , data separator; if empty then space between data
FEXT dat file extension
INTERVAL 4 data logging interval in s
01HR0 sets NO-monitor to local mode
LOOP starts the logging loop
01DO NO-concentration in ppb
01RP0 reaction chamber pressure in hPa
01RT0 monitor's internal temperature in C
01RT1 inlet temperature in C
01RT2 reaction chamber temperature in C
%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%
data-file 'epYYYYMMDD.dat' description:
#N/A
in the 'ecop.ecm' ini-file
- between the data values in col 2 and col 3 there is always a ',' originating
from the monitor; this feature can't be changed
data file column: description
col 1 date and time YYYYDDMMHHMMSS (in UTC)
col 2 measurement cycle index of the monitor (toggles 0/1)
col 3 NO-concentration in ppb
col 4 reaction chamber pressure in hPa
col 5 monitor's internal temperature in C
col 6 inlet temperature in C
col 7 reaction chamber temperature in C

## NO-monpico_torni
LOGGING DATA FROM Pico Tech (ref. to as pico hereafter) Pico ADC-16 data logger
Program 'ecop_log2.exe' in d:\cld780tr_smear is used to log the data.
Start the logging by double-clicking the (short-cut) batch file icon 'NOmonitor'
on the desktop.
- it reads ini-file 'ecop.ecm' in folder ...\cld780tr_smear\data
- data is written to file 'epYYYYMMDD.dat' in folder ...\cld780tr_smear\data
To stop the logging and exit the program press Ctrl-C in the command window, and
confirm.
ini-file 'ecop.ecm' description:
- # means comment line; works only if the first character of a line; don't insert
any comments after the commands
- don't add extra empty lines at the end of the file
%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%
ini-file 'ecop.ecm'
line: description:
# ecop
COMPORT com3 serial port number
#COMPORT \\.\com3
SEPARATOR , data separator; if empty then space between data
FEXT dat file extension
INTERVAL 4 data logging interval in s
01HR0 sets NO-monitor to local mode
LOOP starts the logging loop
01DO NO-concentration in ppb
01RP0 reaction chamber pressure in hPa
01RT0 monitor's internal temperature in C
01RT1 inlet temperature in C
01RT2 reaction chamber temperature in C
%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%%
data-file 'epYYYYMMDD.dat' description:
#N/A
in the 'ecop.ecm' ini-file
- between the data values in col 2 and col 3 there is always a ',' originating
from the monitor; this feature can't be changed
data file column: description
col 1 date and time YYYYDDMMHHMMSS (in UTC)
col 2 measurement cycle index of the monitor (toggles 0/1)
col 3 NO-concentration in ppb
col 4 reaction chamber pressure in hPa
col 5 monitor's internal temperature in C
col 6 inlet temperature in C
col 7 reaction chamber temperature in C

## IONIDMPS
Ionidmps\C:\Ioni2005\RDYYMMDD.DAT
Datafile column | Measured Quantity | Unit in datafile | Description | Notes
1 | Time | hour (decimals)
2 | ?
3 | ?
4 | ?
5 | Particles
6 | DMA Pressure | hPa
7 | DMA Temperature | °C
8 | DMA RH | %

## REACPC
REA\C:\reacpc\RCYYMMDD.DAT
Datafile column | Measured Quantity | Unit in datafile | Description | Notes
1 | Time | hour (decimals)
2 | Up
3 | Dead
4 | Down

## Particle
Hitu\C:\apps\DMPS\data\DMYYMMDD.DAT
mm-dd-yyyy hh:mm:ss'
0.00038 | 4 | 0.025 | 0.033 | 0.109 | 69 | 3025
74.1118 | particles/cm3
40.8738 | particles/cm3
22.5013 | particles/cm3
12.3705 | particles/cm3
6.7943 | particles/cm3
0 | 0
mm-dd-yyyy hh:mm:ss'
0.000175 | 1.01 | 0.025 | 0.033 | 0.28 | 8 | 3010
0 | 0
13.2865 | particles/cm3
18.3205 | particles/cm3
25.2398 | particles/cm3
34.7365 | particles/cm3
47.7479 | particles/cm3
65.5375 | particles/cm3
89.7993 | particles/cm3
122.79 | particles/cm3
167.4922 | particles/cm3
227.8113 | particles/cm3
308.8039 | particles/cm3
416.9324 | particles/cm3
560.3314 | particles/cm3
749.0632 | particles/cm3
995.3418 | particles/cm3
1313.693 | particles/cm3
1721.05 | particles/cm3
2236.8 | particles/cm3
2882.848 | particles/cm3
3683.801 | particles/cm3
4667.385 | particles/cm3
5865.194 | particles/cm3
7313.777 | particles/cm3
9055.986 | particles/cm3
11142.46 | particles/cm3

## Photo config
SMEAR II
PHOTO
Channel number in configuration file | Nokeval + AD-converter configuration | Nokeval channel setting | Measurement | units of measurement | Descriptions | Notes
|  |  | Date& time | yyyymmddhhmmss
|  |  | Moment | s
| AD-converter |  | Installed in Louhi PC
0 | 1 |  | CO2 Photo | ppm |  | Ok 180716
1 | 2 |  | H2O Photo | g/m3 |  | Ok 180716
2 | 3 |  | Used by Mast |  |  | Disconnected
3 | 4 |  | Used by Mast
4 | 5 |  | Used by Mast
5 | 6 |  | Δp URAS | hPa | URAS sample press drop | Ok 180716
6 | 7 |  | MFC supply | ml/min | cuv. Supply air flow | Ok 180716
7 | 8 |  | MFC sample | ml/min | URAS sample flow | Ok 180716
| Nokeval 2 |  | Port 6 in Louhi PC |  |  | Does this Nokeval still exist?
8 | 1 | TCT | T air | ºC
9 | 2 | TCT
10 | 3 | Pt100 | T cuv | ºC
11 | 4 | TCT
12 | 5 | Pt100 | T cuv | ºC
13 | 6 | TCT
14 | 7 | TCT | T
15 | 8 | Pt100 | T cuv | ºC
16 | 9 | TCT
17 | 10 | TCT | T | ºC
18 | 11 | TCT | T
19 | 12 | TCT | T
20 | 13 | +-25 mV
21 | 14 | +-25 mV | PAR cuv | μmol×m-2×s-1
22 | 15 | +-25 mV | PAR cuv | μmol×m-2×s-1
23 | 16 | +-25 mV | PAR cuv | μmol×m-2×s-1
| Nokeval 11 |  | Port 1 in Louhi PC
24 | 1 |  | Pamb | hPa |  | Ok 180716
25 | 2 |  | Used by Mast
26 | 3 |  | Used by Mast
27 | 4 |  | Used by Mast
28 | 5 |  | Photo NO ch. Press. | hPa |  | sensor conn. 25.3.2004
29 | 6 |  | Photo NOx ch. Press. | hPa |  | sensor conn. 25.3.2004, Disabled
30 | 7 |  | Used by Mast |  |  | Disabled
31 | 8 |  | Used by Mast
32 | 9 |  | Used by Mast
33 | 10 |  | NOx Photo | ppb |  | Disabled
34 | 11 |  | Flow (O3 + NO + NOx) | ml×min | O3 + NO + NOx sample flow | Disabled
35 | 12 |  | NO Photo | ppb |  | Disabled
36 | 13 |  | Used by Mast
37 | 14 |  | O3 Photo | ppb | + | Disabled
38 | 15 |  | Δp inlet O3+NOx | hPa |  | Disabled
39 | 16 |  | Used by Mast
| Nokeval 0 |  | Port 6 in Louhi PC
40 | 1 |  | Tcuv Pit 80 | ºC |  | Ok 180716
41 | 2 |  | Tcuv Pit 160 | ºC |  | Ok 180716
42 | 3
43 | 4 |  | Tcuv Pit 70 | ºC |  | Ok 180716
44 | 5
45 | 6
46 | 7
47 | 8
48 | 9
49 | 10
50 | 11
51 | 12
52 | 13
53 | 14
54 | 15
55 | 16
| Nokeval 12 |  | Port 2 in Louhi PC |  |  | Does this Nokeval still exist?
56 | 1 |  | UVA | mV | UVA-sensor in Radiation tower | Disconnected
57 | 2 |  | UVB | mV | UVB-sensor in Radiation tower | Disconnected
58 | 3
59 | 4
60 | 5 |  | PAR | mV | Delta sensor in Radiation Tower | Disconnected?
61 | 6
62 | 7
63 | 8
64 | 9
65 | 10
66 | 11
67 | 12
68 | 13
69 | 14 |  | Global | mV | global radiation | Disconnected?
70 | 15
71 | 16 |  | PAR | mV | PAR sensor in Radiation Tower | Disconnected?
| Nokeval 0 |  | Port 4 in Louhi PC |  |  | Unreliable instrumentation?
72 | 1 |  | PAR cuv | mV
73 | 2
74 | 3 |  | UVB | mV
75 | 4
76 | 5
77 | 6
78 | 7
79 | 8 |  | T ambient | ºC
80 | 9
81 | 10
82 | 11
83 | 12
84 | 13
85 | 14
86 | 15
87 | 16
| Nokeval 1 |  | Port 4 in Louhi PC |  |  | Unreliable instrumentation?
88 | 1
89 | 2
90 | 3
91 | 4
92 | 5
93 | 6
94 | 7
95 | 8
96 | 9
97 | 10
98 | 11
99 | 12
100 | 13
101 | 14
102 | 15
103 | 16
| Nokeval 0 |  | Port 11 in Louhi PC |  | New photo tower
104 | 1 | x | T |  |  | Ok 180716 | 230817 new pt100
105 | 2 |  | PAR |  | Q52959 | Ok 180716 | 230817 PAR
106 | 3 | x | T |  |  | Ok 180716 | 230817 thermocouple
107 | 4 | x | T |  | Long PT100 since spring 2018 | Ok 180716 | 230817 thermocouple
108 | 5 |  | PAR |  | Q104926 | Ok 180716, Long PT100 since spring 2018 until 190218 | 230817 PAR
109 | 6 |  | PAR |  | Serial number? An old one anyway… | Ok 180716 | 230817 long pt100
110 | 7 | x | PAR |  | Q30257 | Ok 180716 | 230817 long pt100
111 | 8 |  | T |  | Long PT100 since spring 2018 | Ok 180716 | 230817 thermocouple
112 | 9 | x | Ambient T |  | Rotronic RS24T ventilated shield | Ok 180716 | 230817 rotronic
113 | 10 |  |  |  |  | Rotronic | 230817 PAR
114 | 11 |  |  |  |  | Rotronic | 230817 new pt100
115 | 12 | x | T |  |  | Ok 180716 | 230817 new pt100
116 | 13 |  |  |  |  | Ok 180716 | 230817 pt100
117 | 14 | x | Universal PAR |  | LiCor 190, serial number? | Ok 180716 | 230817 Univ.PAR
118 | 15 | x | PAR |  | Q104941 without smv-adapter | Ok 180716 | 230817 PAR
119 | 16 |  |  |  |  | Protective earth?

## Photo
SMEAR II |  | Column order in Photo-datafiles
Louhi\D:\smear\photo\SMYYMMDD.sno (sno=shoot number)
Valid 18.04.2002 --> (see Notes for some details)
Datafile column | Measured Quantity | Unit in datafile | Description | Notes
1 | date & time | yyyymmddhhmmss
2 | moment | s | seconds in meas. cycle
3 | T cuv | º C | temperature in cuvette
4 | T air | º C | ambient temperature
5 | PAR cuv | µmolm-2s-1 | PAR radiation | sensor installed on/beside cuvette
6 | UVA | mV | UVA radiation | sensor installed in Sun-tower | In measurements for cuvette 142/144/145 this is PAR above ground.
7 | Flow (poison) | ml/min | poison gas sample flow; | at T=273K & p=sample press.
8 | NO | ppb | NO concentration
9 | Dp (poison) | hPa | poison gas sample press drop
10 | H2O | g/m3 | H2O concentration | AD-conv; at T=273 & p=1013 hPa
11 | MFC supply | ml/min | cuv. supply air flow | T=273 K, p=1013 hpa
12 | CO2 | ppm | CO2 concentration | AD-converter
13 | Dp URAS | hPa | URAS sample press drop
14 | O3 | ppb | O3 concentration
15 | MFC sample | ml/min | URAS sample flow | at T=273 K & p=1013 hPa
16 | Pamb | hPa | ambient pressure
17 | UVB | mV | UVB radiation | sensor installed in Sun-tower | In measurements for cuvette 142/144/145/153 this is PAR above ground.
18 | NOx | ppb | NOx concentration |  | In measurements for cuvette 142/144/145/153 this is PAR above ground.
19 | NO chamber press | hPa |  | Since 21.05.2010
20 | NOx chamber press | hPa |  | Since 21.05.2010
Several other version have been used between 2002 and 2018!
Valid 21.03.2018 --> (see Notes for some details), but likely missing one channel (check the next version)!
Datafile column | Measured Quantity | Unit in datafile | Description | Notes
1 | date & time | yyyymmddhhmmss
2 | moment | s | seconds in meas. cycle
3 | T cuv | º C | temperature in cuvette
4 | T air | º C | ambient temperature
5 | PAR cuv | µmolm-2s-1 | PAR radiation | sensor installed on/beside cuvette
6 | - | ?? | Nothing | Random signal from some working sensor
7 | - | ?? | Nothingg | Random signal from some working sensor
8 | - | ?? | Nothing | Random signal from some working sensor
9 | - | ?? | Nothing | Random signal from some working sensor
10 | H2O | g/m3 | H2O concentration | AD-conv; at T=273 & p=1013 hPa
11 | MFC supply | ml/min | cuv. supply air flow | T=273 K, p=1013 hpa
12 | CO2 | ppm | CO2 concentration | AD-converter
13 | Dp Licor | hPa | Licor sample press drop
14 | - | ?? | Nothing | Random signal from some working sensor
15 | MFC sample | ml/min | Licor sample flow | at T=273 K & p=1013 hPa
16 | Pamb | hPa | ambient pressure
17 | - | ?? | Nothing | Random signal from some working sensor
18 | PAR univ | µmolm-2s-1 | PAR radiation | Sensor installed in Photo 2 tower
19 | - | ?? | Nothing | Random signal from some working sensor
20 | - | ?? | Nothing | Random signal from some working sensor
Valid 28.03.2019 --> (see Notes for some details); AND LIKELY VALID ALSO STARTING FROM 21.3.2018
Datafile column | Measured Quantity | Unit in datafile | Description | Notes
1 | date & time | yyyymmddhhmmss
2 | moment | s | seconds in meas. cycle
3 | T cuv | º C | temperature in cuvette
4 | T air | º C | ambient temperature
5 | PAR cuv | µmolm-2s-1 | PAR radiation | sensor installed on/beside cuvette
6 | - | ?? | Nothing | Random signal from some working sensor
7 | - | ?? | Nothingg | Random signal from some working sensor
8 | - | ?? | Nothing | Random signal from some working sensor
9 | - | ?? | Nothing | Random signal from some working sensor
10 | H2O | g/m3 | H2O concentration | AD-conv; at T=273 & p=1013 hPa
11 | MFC supply | ml/min | cuv. supply air flow | T=273 K, p=1013 hpa
12 | CO2 | ppm | CO2 concentration | AD-converter
13 | Dp Licor | hPa | Licor sample press drop
14 | - | ?? | Nothing | Random signal from some working sensor
15 | MFC sample | ml/min | Licor sample flow | at T=273 K & p=1013 hPa
16 | Pamb | hPa | ambient pressure
17 | - | ?? | Nothing | Random signal from some working sensor
18 | - | ?? | Nothing | Random signal from some working sensor
19 | PAR univ | µmolm-2s-1 | PAR radiation | Sensor installed in Photo 2 tower
20 | - | ?? | Nothing | Random signal from some working sensor
21 | - | ?? | Nothing | Random signal from some working sensor

## Rain
Louhi\D:\smear\rain\SMYYMMDD.RAI |  |  |  | Began 24.8.1996
Datafile column | Nokeval12 Channel | Measured Quantity | Unit in datafile | Notes
1 |  | date and time | yyyymmddhhmmss
2 | 7 | tipping bucket counter | rain mm | ARG-100 in tower
3 | 8 | SWS | mV | Surface Wetness Sensor

## WTESTI
WTEST |  | CO2 soil profiles at Pit 100, began 9.6.2004 |  |  |  | WTEST |  | CO2 soil profiles at Pit 100, modified 1.4.2007 to Liisa's Vaisala chamber |  |  | WTEST |  | CO2 soil profiles at Pit 100, modified 23.4.2008 to Liisa's Vaisala chamber
Louhi\D:\smear\SOIL_CO2\SMYYMMDD.WAI | [5 empty] | Louhi\D:\smear\SOIL_CO2\SMYYMMDD.WAI |  |  |  |  | Louhi\D:\smear\SOIL_CO2\SMYYMMDD.WAI
[6 empty] | Soil CO2 underground & chamber at Pit 100 |  |  |  |  | Soil CO2 underground & chamber at Pit 100
Datafile column | Nokeval 8 channel | label | units | Notes |  | Datafile column | Nokeval 8 channel | label | units | Notes | Datafile column | Nokeval 8 channel | label | units | Notes
1 |  | date and time | yyyymmddhhss |  |  | 1 |  | date and time | yyyymmddhhss | 1.4.2007 - 22.04.2008 | 1 |  | date and time | yyyymmddhhss | 23.04.2008 - 20.04.2009
2 | 1 | Heat sensor | °C |  |  | 2 | 1 | Heat sensor | °C |  | 2 | 1 | Heat sensor | °C
3 | 2 | Heat sensor | °C |  |  | 3 | 2 | Heat sensor | °C |  | 3 | 2 | Heat sensor | °C
4 | 3 | Heat sensor | °C |  |  | 4 | 3 | Heat sensor | °C |  | 4 | 3 | Heat sensor | °C
5 | 4 | Heat sensor | °C |  |  | 5 | 4 | Heat sensor | °C |  | 5 | 4 | Heat sensor | °C
6 | 5 | Temperature 22 cm | °C |  |  | 6 | 5 | Temperature 22 cm | °C |  | 6 | 5 | Temperature 22 cm | °C
7 | 6 | Temperature 12 cm | °C |  |  | 7 | 6 | Temperature 12 cm | °C |  | 7 | 6 | Temperature 12 cm | °C
8 | 7 | Temperature 0 cm | °C |  |  | 8 | 7 | PAR in cuvette |  |  | 8 | 7 | Chamber open/closed |  | 690mV=closed, 1137mV=open
9 | 8 | Temperature humus | °C |  |  | 9 | 8 | PAR in cuvette |  |  | 9 | 8 | Temperature in chamber | °C
10 | 9 | CO2 0 cm | ppm/5 |  |  | 10 | 9 | PAR in cuvette |  |  | 10 | 9 | CO2 in chamber | ppm
11 | 10 | CO2 humus | ppm/5 |  |  | 11 | 10 | Temperature in cuvette | °C |  | 11 | 10 | PAR in chamber | mV
12 | 11 | CO2 12 cm | ppm/5 |  |  | 12 | 11 | Chamber open/closed |  | 690mV=closed, 1137mV=open | 12 | 11 | Empty
13 | 12 | CO2 22 cm | ppm/5 |  |  | 13 | 12 | CO2 in chamber | ppm |  | 13 | 12 | Empty
14 | 13 | CO2 GMM-220 A 16 cm | ppm/10 |  |  | 14 | 13 | CO2 0 cm | ppm/5 |  | 14 | 13 | CO2 0 cm | ppm/5
15 | 14 | CO2 GMM-220 B 16 cm | ppm/10 |  |  | 15 | 14 | CO2 humus | ppm/5 |  | 15 | 14 | CO2 humus | ppm/5
16 | 15 | Empty |  |  |  | 16 | 15 | CO2 12 cm | ppm/5 |  | 16 | 15 | CO2 12 cm | ppm/5
17 | 16 | Empty |  |  |  | 17 | 16 | CO2 22 cm | ppm/5 |  | 17 | 16 | CO2 22 cm | ppm/5
Louhi\D:\smear\SOIL_CO2\SMYYMMDD.WAI
Soil CO2 underground & ToHeLi chambers 1 and 2 at Pit 100, between 13.05.2009 -
Datafile column | Nokeval 8 channel | label | units | Notes
1 |  | datetime | yyyymmddhhss | Began 13.05.2009
2 | 1 | Heat sensor | °C
3 | 2 | Heat sensor | °C
4 | 3 | Heat sensor | °C
5 | 4 | Heat sensor | °C
6 | 5 | Temperature 22 cm | °C
7 | 6 | Temperature 12 cm | °C |  | 13.08. 2010 NOTICED BY JP
8 | 7 | CO2 humus | mV | ppm=(4500/425.19)*Mv<-WRONG | , CORRECT PPM=10.759*Mv
9 | 8 | CO2 2cm | mV | ppm=(4500/251.99)*mV<-WRONG | , CORRECT PPM=10.956*Mv
10 | 9 | CO2 12cm | mV | ppm=(4500/424.9)*mV<-WRONG | , CORRECT PPM=22.961*Mv
11 | 10 | CO2 22cm | mV | ppm=(4500/206.16)*mV<-WRONG | , CORRECT PPM=23.673*Mv
12 | 11 | T | °C
13 | 12 | T | °C
14 | 13 | T | °C
15 | 14 | T | °C
16 | 15 | Chamber1 CO2 | ppm
17 | 16 | Chamber1 open/closed | 270/4.8 mV
Nok 9
18 | 1 | Chamber1 T air | °C
19 | 2 | Chamber1 T ground | °C
20 | 3 | Chamber1 PAR 1 | mV
21 | 4 | Chamber1 PAR 2 | mV
22 | 5 | Chamber1 RH | mV
23 | 6 | Chamber1 T (RH)
24 | 7 | Chamber2 CO2 | ppm
25 | 8 | Chamber2 open/closed | 1222/5 mV
26 | 9 | Chamber2 T air | °C
27 | 10 | Chamber2 T ground | °C
28 | 11 | Chamber2 PAR 1 | mV
29 | 12 | Chamber2 PAR 2 | mV
30 | 13 | Chamber2 RH | mV
31 | 14 | Chamber2 T (RH)
32 | 15 | Empty
33 | 16 | Empty

## CO2_p80
CO2_p80 |  | CO2 soil profiles at Pit 80, began 6.12.2005
Sampo\D:\smear\CO2_p80\SMYYMMDD.COP
05.12.2005 - 18.05.2006
Datafile column | Nokeval 0 channel | Label | Units | Notes
1 |  | date and time | yyyymmhhss
2 | 1 | CO2 -22cm |  | GMM222
3 | 2 | CO2 -12cm |  | GMP343
4 | 3 | CO2 -7cm |  | GMP343
5 | 4 | CO2 -2cm |  | GMP343
6 | 5 | CO2 +5cm |  | GMP343
7 | 6 | CO2 +15cm |  | GMP343
8 | 7 | CO2 +30cm |  | GMP343
9 | 8 | CO2 +1m |  | GMP343
10 | 9 | CO2 +1m |  | GMP343
11 | 10 | T
12 | 11 | T
13 | 12 | T
14 | 13 | T
15 | 14 | T
16 | 15 | T
17 | 16 | T
Nokeval 1
18 | 1 | T
19 | 2 | T
20 | 3 | T instrument box
21 | 4 | empty
22 | 5 | empty
23 | 6 | empty
24 | 7 | empty
25 | 8 | empty
26 | 9 | LiCor PAR
27 | 10 | LiCor PAR
28 | 11 | Delta PAR
29 | 12 | Delta PAR
30 | 13 | Fotodiode
31 | 14 | Fotodiode
32 | 15 | Fotodiode
33 | 16 | Fotodiode
Sampo\D:\smear\CO2_p80\SMYYMMDD.COP
19.05.2006 - 09.05.2007
Datafile column | Nokeval channel | Label | Units | Notes
1 |  | date and time | yyyymmhhss
2 | 1 | CO2 -22cm, C |  | GMM222, profile
3 | 2 | CO2 -12cm, B |  | GMP343, profile
4 | 3 | CO2 -7cm, A |  | GMP343, profiili
5 | 4 | CO2 -2cm, H |  | GMP343, profiili
6 | 5 | CO2, C |  | GMP343, girdling
7 | 6 | CO2, B |  | GMP343, girdling
8 | 7 | CO2, A |  | GMP343, girdling
9 | 8 | CO2, H |  | GMP343, girdling
10 | 9 | CO2 |  | GMP343, girdling
11 | 10 | T
12 | 11 | T
13 | 12 | T
14 | 13 | T
15 | 14 | T
16 | 15 | T, -20 cm |  | girdling
17 | 16 | T, -10 cm |  | girdling
Nokeval 1
18 | 1 | T, -5 cm |  | girdling
19 | 2 | T
20 | 3 | T instrument box
21 | 4 | empty
22 | 5 | empty
23 | 6 | empty
24 | 7 | empty
25 | 8 | empty
26 | 9 | empty
27 | 10 | empty
28 | 11 | empty
29 | 12 | empty
30 | 13 | empty
31 | 14 | empty
32 | 15 | empty
33 | 16 | empty
Sampo\D:\smear\CO2_p80\SMYYMMDD.COP
10.05.2007 - 30.05.2008
Datafile column | Nokeval 0 channel | Label | Units | Notes
1 |  | date and time | yyyymmhhss
2 | 1 | CO2 -22cm |  | GMM222
3 | 2 | CO2 -12cm |  | GMP343
4 | 3 | CO2 -7cm |  | GMP343
5 | 4 | CO2 -2cm |  | GMP343
6 | 5 | CO2 in chamber |  | GMP343
7 | 6 | RH in chamber
8 | 7 | Empty
9 | 8 | Chamber open/closed
10 | 9 | CO2 ambient air |  | GMP343
11 | 10 | T
12 | 11 | T
13 | 12 | T
14 | 13 | T
15 | 14 | T
16 | 15 | T
17 | 16 | T
Nokeval 1
18 | 1 | T
19 | 2 | T
20 | 3 | T instrument box
21 | 4 | Empty
22 | 5 | Empty
23 | 6 | Empty
24 | 7 | Empty
25 | 8 | Empty
26 | 9 | Empty
27 | 10 | Empty
28 | 11 | Empty
29 | 12 | Empty
30 | 13 | T chamber
31 | 14 | T chamber
32 | 15 | Empty
33 | 16 | Empty
Sampo\D:\smear\CO2_p80\SMYYMMDD.COP
31.05.2008 - 12.05.2009
Datafile column | Nokeval 0 channel | Label | Units | Notes
1 |  | date and time | yyyymmhhss
2 | 1 | CO2 -22cm |  | GMM222
3 | 2 | CO2 -12cm |  | GMP343
4 | 3 | CO2 -7cm |  | GMP343
5 | 4 | CO2 -2cm |  | GMP343
6 | 5 | CO2 in chamber |  | GMP343
7 | 6 | RH in chamber
8 | 7 | Empty
9 | 8 | Chamber open/closed
10 | 9 | CO2 ambient air |  | GMP343
11 | 10 | T
12 | 11 | T
13 | 12 | T
14 | 13 | T
15 | 14 | T
16 | 15 | T
17 | 16 | T
Nokeval 1
18 | 1 | T
19 | 2 | T
20 | 3 | T instrument box
21 | 4 | Empty
22 | 5 | Empty
23 | 6 | Empty
24 | 7 | Empty
25 | 8 | Empty
26 | 9 | Empty
27 | 10 | Empty
28 | 11 | Empty
29 | 12 | Empty
30 | 13 | T chamber
31 | 14 | T chamber
32 | 15 | Empty
33 | 16 | Empty
Sampo\D:\smear\CO2_p80\SMYYMMDD.COP
18.05.2009 -
Datafile column | Nokeval 0 channel | Label | Units | Notes
1 |  | date and time | yyyymmhhss
2 | 1 | CO2 -22cm |  | GMM222
3 | 2 | CO2 -12cm |  | GMP343 | THIS SENSOR (SRNO A3330008_0-5000 PPM) WAS TEMPORARILY REMOVED FOR LAKE PROFILE ON 27.5.2010.
4 | 3 | CO2 -7cm |  | GMP343
5 | 4 | CO2 -2cm |  | GMP343
6 | 5 | Empty
7 | 6 | Empty
8 | 7 | Empty
9 | 8 | Empty
10 | 9 | CO2 ambient air |  | GMP343
11 | 10 | T10
12 | 11 | T9
13 | 12 | T8
14 | 13 | T7
15 | 14 | T6
16 | 15 | Empty
17 | 16 | Empty
Nokeval 1
18 | 1 | Empty
19 | 2 | T2
20 | 3 | T1 instrument box
21 | 4 | Empty
22 | 5 | Empty
23 | 6 | Empty
24 | 7 | Empty
25 | 8 | Empty
26 | 9 | Chamber3 CO2 | ppm
27 | 10 | Chamber3 open/closed | 272/5.9 mV
28 | 11 | Chamber3 T air | °C
29 | 12 | Chamber3 T ground | °C
30 | 13 | Chamber3 PAR 1 | mV
31 | 14 | Chamber3 PAR 2 | mV
32 | 15 | Chamber3 RH | mV
33 | 16 | Chamber3 T (RH)

## Old_Soil
File Smyymmdd.maa |  |  |  | 5.8.1995 - 30.5.2001
Column | Com |  | Channel | Label | depth of | Length of | Unit | Notes | Conversion
[5 empty] | sensor, cm | tensionmeters
| COM 5 | Noke |  |  |  | tube / termistor no.
1 |  |  |  | date and time |  |  | yyyymmddhhmmss
2 |  | 0 | 1 | P 50 A | 4 | 36 cm | mV
3 |  |  | 2 | P 50 B1 | 17 | 36 cm | mV
4 |  |  | 3 | P 50 B2 | 27 | 36 cm | mV
5 |  |  | 4 | T in cuvette ?
6 |  |  | 5 | T in cuvette ?
7 |  |  | 6 | P 160 A | 4 | 36 cm | mV
8 |  |  | 7 | P 160 B1 | 17 | 36 cm | mV
9 |  |  | 8 | P 160 B2 | 33 | 66 cm | mV
10 |  |  | 9 | P 160 C1 | 53 | 66 cm | mV
11 |  |  | 10 | P 160 C2 | 93 | 126 cm | mV
12 |  |  | 11 | P 160 C3 | 110 | 126 cm | mV
13 |  |  | 12 | T 50 AIR |  | KTY951 | °C
14 |  |  | 13 | T 50 H |  | KTY952 | °C
15 |  |  | 14 | T 50 A | 5 | KTY953 | °C
16 |  |  | 15 | T 50 B | 11 | KTY954 | °C
17 |  |  | 16 | T 50 C | 22 | KTY955 | °C
| COM 5 | Noke | ch | label
18 | 17 | 1 | 1 | T 130 H/A | 0 | KTY9517 | °C
19 | 18 |  | 2 | T 130 A | 5 | KTY9518 | °C
20 | 19 |  | 3 | T 130 B1 | 9 | KTY9519 | °C
21 | 20 |  | 4 | T 130 B2 | 23 | KTY9520 | °C
22 | 21 |  | 5 | T 130 C1 | 58 | KTY9521 | °C
23 | 22 |  | 6 | T 130 C2 | 90 | KTY9522 | °C
24 | 23 |  | 7 | T 130 B | 32 | KTY9523 | °C | STONE
25 | 24 |  | 8 | T 130 C | 102 | KTY9524 | °C | BEDROCK
26 | 25 |  | 9 | T 160 H |  | KTY9525 | °C
27 | 26 |  | 10 | T 160 A | 3 | KTY9526 | °C
28 | 27 |  | 11 | T 160 B1 | 14 | KTY9527 | °C
29 | 28 |  | 12 | T 160 B2 | 29 | KTY9528 | °C
30 | 29 |  | 13 | T 160 C1 | 45 | KTY9529 | °C
31 | 30 |  | 14 | T 160 C2 | 87 | KTY9530 | °C
32 | 31 |  | 15 | T 160 C3 | 125 | KTY9531 | °C
33 | 32 |  | 16 | T 160 AIR |  | KTY9532 | °C
| COM 5 | Noke | ch | label
34 | 33 | 2 | 1 | P 100 A | 4 | 36 cm | mV
35 | 34 |  | 2 | P 100 B | 16 | 36 cm | mV
36 | 35 |  | 3 | P 100 C1 | 31 | 66 cm | mV
37 | 36 |  | 4 | P 100 C2 | 51 | 66 cm | mV
38 | 37 |  | 5 | P 100 C3 | 70 | 96 cm | mV
39 | 38 |  | 6 | T 70 H/A | 0 | KTY956 | °C
40 | 39 |  | 7 | T 70 B1 | 5 | KTY957 | °C
41 | 40 |  | 8 | T 70 B2 | 11 | KTY958 | °C
42 | 41 |  | 9 | T 70 B3 | 23 | KTY959 | °C
43 | 42 |  | 10 | T 70 C1 | 55 | KTY9510 | °C
44 | 43 |  | 11 | T 100 H |  | KTY9511 | °C
45 | 44 |  | 12 | T 100 A | 2 | KTY9512 | °C
46 | 45 |  | 13 | T 100 B1 | 10 | KTY9513 | °C
47 | 46 |  | 14 | T 100 B2 | 22 | KTY9514 | °C
48 | 47 |  | 15 | T 100 C1 | 42 | KTY9515 | °C
49 | 48 |  | 16 | T 100 C2 | 66 | KTY9516 | °C

## Soil 30-05-2001
File Smyymmdd.maa |  |  | Valid until 7. June 2000 | [5 empty] | File Smyymmdd.maa |  |  | Valid since 31. May. 2001
Column | Com | Channel | Label | depth of | Length of | Unit | Notes | Conversion | Column | Com | Channel | Label | depth of | Length of | Unit | Notes | Conversion
|  |  |  | sensor, cm | tensionmeters | [7 empty] | sensor, cm | tensionmeters
| COM 5 | Noke 0 |  |  | tube / termistor no. |  |  |  |  | COM 5 | Noke 0 |  |  | tube / termistor no.
1 |  |  | date and time |  |  | yyyymmddhhmmss |  |  | 1 |  |  | date and time |  |  | yyyymmddhhmmss
2 |  | 1 | P 50 A | 4 | 36 cm | mV |  |  | 2 |  | 1 | P 160 B2 | 33 | 66 cm | mV
3 |  | 2 | P 50 B1 | 17 | 36 cm | mV |  |  | 3 |  | 2 | Empty
4 |  | 3 | P 50 B2 | 27 | 36 cm | mV |  |  | 4 |  | 3 | P 50 B2 | 27 | 36 cm | mV
5 |  | 4 | T in cuvette ? | [5 empty] | 5 |  | 4 | T in cuvette PIT 70 |  |  | °C | Cuvette 10
6 |  | 5 | T in cuvette ? | [5 empty] | 6 |  | 5 | P 50 A | 4 | 36 cm | mV
7 |  | 6 | P 160 A | 4 | 36 cm | mV |  |  | 7 |  | 6 | P 50 B1 | 17 | 36 cm | mV
8 |  | 7 | P 160 B1 | 17 | 36 cm | mV |  |  | 8 |  | 7 | P 160 A | 4 | 36 cm | mV
9 |  | 8 | P 160 B2 | 33 | 66 cm | mV |  |  | 9 |  | 8 | P 160 B1 | 17 | 36 cm | mV
10 |  | 9 | P 160 C1 | 53 | 66 cm | mV |  |  | 10 |  | 9 | P 160 C1 | 53 | 66 cm | mV
11 |  | 10 | P 160 C2 | 93 | 126 cm | mV |  |  | 11 |  | 10 | P 160 C2 | 93 | 126 cm | mV
12 |  | 11 | P 160 C3 | 110 | 126 cm | mV |  |  | 12 |  | 11 | P 160 C3 | 110 | 126 cm | mV
13 |  | 12 | T 50 AIR |  | KTY951 | °C |  |  | 13 |  | 12 | T in cuvette PIT 80 |  |  | °C | Cuvette 11
14 |  | 13 | T 50 H |  | KTY952 | °C |  |  | 14 |  | 13 | T 50 H |  | KTY952 | °C
15 |  | 14 | T 50 A | 5 | KTY953 | °C |  |  | 15 |  | 14 | T 50 A | 5 | KTY953 | °C
16 |  | 15 | T 50 B | 11 | KTY954 | °C |  |  | 16 |  | 15 | T 50 B | 11 | KTY954 | °C
17 |  | 16 | T 50 C | 22 | KTY955 | °C |  |  | 17 |  | 16 | T 50 C | 22 | KTY955 | °C
| COM 5 | Noke 1 | label | [6 empty] | COM 5 | Noke 1 | label
18 | 17 | 1 | T 130 H/A | 0 | KTY9517 | °C |  |  | 18 | 17 | 1 | T 130 H/A | 0 | KTY9517 | °C
19 | 18 | 2 | T 130 A | 5 | KTY9518 | °C |  |  | 19 | 18 | 2 | T 130 A | 5 | KTY9518 | °C
20 | 19 | 3 | T 130 B1 | 9 | KTY9519 | °C |  |  | 20 | 19 | 3 | T 130 B1 | 9 | KTY9519 | °C
21 | 20 | 4 | T 130 B2 | 23 | KTY9520 | °C |  |  | 21 | 20 | 4 | T 130 B2 | 23 | KTY9520 | °C
22 | 21 | 5 | T 130 C1 | 58 | KTY9521 | °C |  |  | 22 | 21 | 5 | T 130 C1 | 58 | KTY9521 | °C
23 | 22 | 6 | T 130 C2 | 90 | KTY9522 | °C |  |  | 23 | 22 | 6 | T 130 C2 | 90 | KTY9522 | °C
24 | 23 | 7 | T 130 B | 32 | KTY9523 | °C | STONE |  | 24 | 23 | 7 | T 130 B | 32 | KTY9523 | °C | STONE
25 | 24 | 8 | T 130 C | 102 | KTY9524 | °C | BEDROCK |  | 25 | 24 | 8 | T 130 C | 102 | KTY9524 | °C | BEDROCK
26 | 25 | 9 | T 160 A | 3 | KTY9526 | °C | <-these have been mixed in 1997 |  | 26 | 25 | 9 | T 160 A | 3 | KTY9526 | °C
27 | 26 | 10 | T 160 B1 | 14 | KTY9527 | °C | see the diary |  | 27 | 26 | 10 | T 160 B1 | 14 | KTY9527 | °C
28 | 27 | 11 | T 160 B2 | 29 | KTY9528 | °C | <-these have been mixed in 1997 |  | 28 | 27 | 11 | T 160 B2 | 29 | KTY9528 | °C
29 | 28 | 12 | T 160 H |  | KTY9525 | °C | see the diary |  | 29 | 28 | 12 | T 160 H |  | KTY9525 | °C
30 | 29 | 13 | T 160 C1 | 45 | KTY9529 | °C |  |  | 30 | 29 | 13 | T 160 C1 | 45 | KTY9529 | °C
31 | 30 | 14 | T 160 C2 | 87 | KTY9530 | °C |  |  | 31 | 30 | 14 | T 160 C2 | 87 | KTY9530 | °C
32 | 31 | 15 | T 160 C3 | 125 | KTY9531 | °C |  |  | 32 | 31 | 15 | T 160 C3 | 125 | KTY9531 | °C
33 | 32 | 16 | T 160 AIR |  | KTY9532 | °C |  |  | 33 | 32 | 16 | T 160 AIR |  | KTY9532 | °C
| COM 5 | Noke 2 | label | [6 empty] | COM 5 | Noke 2 | label
34 | 33 | 1 | P 100 A | 4 | 36 cm | mV |  |  | 34 | 33 | 1 | P 100 A | 4 | 36 cm | mV
35 | 34 | 2 | P 100 B | 16 | 36 cm | mV |  |  | 35 | 34 | 2 | P 100 B | 16 | 36 cm | mV
36 | 35 | 3 | P 100 C1 | 31 | 66 cm | mV |  |  | 36 | 35 | 3 | P 100 C1 | 31 | 66 cm | mV
37 | 36 | 4 | P 100 C2 | 51 | 66 cm | mV |  |  | 37 | 36 | 4 | P 100 C2 | 51 | 66 cm | mV
38 | 37 | 5 | P 100 C3 | 70 | 96 cm | mV |  |  | 38 | 37 | 5 | P 100 C3 | 70 | 96 cm | mV
39 | 38 | 6 | T 70 H/A | 0 | KTY956 | °C |  |  | 39 | 38 | 6 | T 70 H/A | 0 | KTY956 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
40 | 39 | 7 | T 70 B1 | 5 | KTY957 | °C |  |  | 40 | 39 | 7 | T 70 B1 | 5 | KTY957 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
41 | 40 | 8 | T 70 B2 | 11 | KTY958 | °C |  |  | 41 | 40 | 8 | T 70 B2 | 11 | KTY958 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
42 | 41 | 9 | T 70 B3 | 23 | KTY959 | °C |  |  | 42 | 41 | 9 | T 70 B3 | 23 | KTY959 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
43 | 42 | 10 | T 70 C1 | 55 | KTY9510 | °C |  |  | 43 | 42 | 10 | T 70 C1 | 55 | KTY9510 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
44 | 43 | 11 | T 100 H |  | KTY9511 | °C |  |  | 44 | 43 | 11 | T 100 H |  | KTY9511 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
45 | 44 | 12 | T 100 A | 2 | KTY9512 | °C |  |  | 45 | 44 | 12 | T 100 A | 2 | KTY9512 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
46 | 45 | 13 | T 100 B1 | 10 | KTY9513 | °C |  |  | 46 | 45 | 13 | T 100 B1 | 10 | KTY9513 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
47 | 46 | 14 | T 100 B2 | 22 | KTY9514 | °C |  |  | 47 | 46 | 14 | T 100 B2 | 22 | KTY9514 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
48 | 47 | 15 | T 100 C1 | 42 | KTY9515 | °C |  |  | 48 | 47 | 15 | T 100 C1 | 42 | KTY9515 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
49 | 48 | 16 | T 100 C2 | 66 | KTY9516 | °C |  |  | 49 | 48 | 16 | T 100 C2 | 66 | KTY9516 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
| COM 2 | Noke | [7 empty] | COM 2 | Noke 4 | Label
50 | [8 empty] | 50 |  | 1 | Tensionmeter | 4 |  |  | Added to Smyymmdd.maa file at 29. May 2001
51 | [8 empty] | 51 |  | 2 | Tensionmeter | 15 |  |  | Added to Smyymmdd.maa file at 29. May 2001
52 | [8 empty] | 52 |  | 3 | Tensionmeter | 4 |  |  | Added to Smyymmdd.maa file at 29. May 2001
53 | [8 empty] | 53 |  | 4 | Tensionmeter | 15 |  |  | Added to Smyymmdd.maa file at 29. May 2001
54 | [8 empty] | 54 |  | 5 | Tensionmeter | 30 |  |  | Added to Smyymmdd.maa file at 29. May 2001
55 | [8 empty] | 55 |  | 6 | Tensionmeter | 30 | NEGATIVE SIGNAL |  | Added to Smyymmdd.maa file at 29. May 2001
56 | [8 empty] | 56 |  | 7 | Tensionmeter | 4 |  |  | Added to Smyymmdd.maa file at 29. May 2001
57 | [8 empty] | 57 |  | 8 | Tensionmeter | 15 | Tensiometer broken 1. June 2001 |  | Added to Smyymmdd.maa file at 29. May 2001
58 | [8 empty] | 58 |  | 9 | Tensionmeter | 27 | Tensiometer broken 1. June 2001 |  | Added to Smyymmdd.maa file at 29. May 2001
59 | [8 empty] | 59 |  | 10 | Tensionmeter | 50 |  |  | Added to Smyymmdd.maa file at 29. May 2001
60 | [8 empty] | 60 |  | 11 | Tensionmeter | 4 |  |  | Added to Smyymmdd.maa file at 29. May 2001
61 | [8 empty] | 61 |  | 12 | Tensionmeter | 15 |  |  | Added to Smyymmdd.maa file at 29. May 2001
62 | [8 empty] | 62 |  | 13 | Tensionmeter | 30 | Tensiometer broken 1. June 2001 |  | Added to Smyymmdd.maa file at 29. May 2001
63 | [8 empty] | 63 |  | 14 | Tensionmeter | 50 | NEGATIVE SIGNAL |  | Added to Smyymmdd.maa file at 29. May 2001
64 | [8 empty] | 64 |  | 15 | Tensionmeter | 4 |  |  | Added to Smyymmdd.maa file at 29. May 2001
65 | [8 empty] | 65 |  | 16 | Tensionmeter | 5 | NEGATIVE SIGNAL |  | Added to Smyymmdd.maa file at 29. May 2001

## Soil 01-08-2005
Louhi\D:\smear\soil\SMYYMMDD.maa |  |  |  | Valid since 15.05.2006
Column | Com | Channel | Label | depth of | Length of | Unit | Notes | Conversion
|  |  |  | sensor, cm | tensionmeters
| COM 5 | Noke 0 |  |  | tube / termistor no.
1 |  |  | date and time |  |  | yyyymmddhhmmss
2 |  | 1 | T in cuvette PIT 80 |  |  | °C | Cuvette 13/14/23, since 08.05.2006
3 |  | 2 | T in cuvette PIT 160 |  |  | °C | Cuvette 25
4 |  | 3 | PAR |  |  | mV | Ending 20.11.2007
5 |  | 4 | T in cuvette PIT 70 |  |  | °C | Cuvette 10/20
6 |  | 5 | Pit 50 27 cm | 27 | 66 | mV | Tensiometer
7 |  | 6 | Pit 50 17 cm | 17 | 36 | mV | Tensiometer
8 |  | 7 | Pit 50 4 cm | 4 | 36 | mV | Tensiometer
9 |  | 8 | Empty
10 |  | 9 | Pit 70 16/b cm sarake10 | 16 | 33 | mV | Tensiometer
11 |  | 10 | Pit 70 4 cm sarake 11 | 4 | 33 | mV | Tensiometer
12 |  | 11 | PAR |  |  | mV | Ending 20.11.2007
13 |  | 12 | Empty |  |  |  | T in cuvette PIT 80, until 08.05.2006
14 |  | 13 | T 50 H |  | KTY952 | °C
15 |  | 14 | T 50 A | 5 | KTY953 | °C
16 |  | 15 | T 50 B | 11 | KTY954 | °C
17 |  | 16 | T 50 C | 22 | KTY955 | °C
| COM 5 | Noke 1 | label
18 | 17 | 1 | T 130 H/A | 0 | KTY9517 | °C
19 | 18 | 2 | T 130 A | 5 | KTY9518 | °C
20 | 19 | 3 | T 130 B1 | 9 | KTY9519 | °C
21 | 20 | 4 | T 130 B2 | 23 | KTY9520 | °C
22 | 21 | 5 | T 130 C1 | 58 | KTY9521 | °C
23 | 22 | 6 | T 130 C2 | 90 | KTY9522 | °C
24 | 23 | 7 | T 130 B | 32 | KTY9523 | °C | STONE
25 | 24 | 8 | T 130 C | 102 | KTY9524 | °C | BEDROCK
26 | 25 | 9 | T 160 A | 3 | KTY9526 | °C
27 | 26 | 10 | T 160 B1 | 14 | KTY9527 | °C
28 | 27 | 11 | T 160 B2 | 29 | KTY9528 | °C
29 | 28 | 12 | T 160 H |  | KTY9525 | °C
30 | 29 | 13 | T 160 C1 | 45 | KTY9529 | °C
31 | 30 | 14 | T 160 C2 | 87 | KTY9530 | °C
32 | 31 | 15 | T 160 C3 | 125 | KTY9531 | °C
33 | 32 | 16 | T 160 AIR |  | KTY9532 | °C
| COM 5 | Noke 2 | label
34 | 33 | 1 | PARtst, ending 20.11.2007 |  |  | mV | Since 01.06.09 MaSa chamber open(0), closed(500), closed-sample(900)
35 | 34 | 2 | PARtst |  |  | mV | Ending 30.05.2008
36 | 35 | 3 | PARtst |  |  | mV | Ending 30.05.2008
37 | 36 | 4 | PARtst |  |  | mV | Ending 30.05.2008
38 | 37 | 5 | T in cuvette PIT 70, humus |  | KTY951 | °C | since 051214
39 | 38 | 6 | T 70 H/A | 0 | KTY956 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
40 | 39 | 7 | T 70 B1 | 5 | KTY957 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
41 | 40 | 8 | T 70 B2 | 11 | KTY958 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
42 | 41 | 9 | T 70 B3 | 23 | KTY959 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
43 | 42 | 10 | T 70 C1 | 55 | KTY9510 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
44 | 43 | 11 | T 100 H |  | KTY9511 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
45 | 44 | 12 | T 100 A | 2 | KTY9512 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
46 | 45 | 13 | T 100 B1 | 10 | KTY9513 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
47 | 46 | 14 | T 100 B2 | 22 | KTY9514 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
48 | 47 | 15 | T 100 C1 | 42 | KTY9515 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
49 | 48 | 16 | T 100 C2 | 66 | KTY9516 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
| COM 2 | Noke 3 | Label
50 |  | 1 | Theta probe, MASA | 50 |  | mV | Since 01.02.07
51 |  | 2 | Theta probe, MASA | 25 |  | mV | Since 01.02.07
52 |  | 3 | Theta probe, MASA | 10 |  | mV | Since 01.02.07
53 |  | 4 | Theta probe, MASA | H |  | mV | Since 01.02.07
54 |  | 5 | T, MASA | 50 |  | °C | Since 20.10.06
55 |  | 6 | T, MASA | 25 |  | °C | Since 20.10.06
56 |  | 7 | T, MASA | 10 |  | °C | Since 20.10.06
57 |  | 8 | T, MASA | H |  | °C | Since 20.10.06
58 |  | 9 | Pit 160 33 cm sarake 58 | 33 | 66 | mV | Tensiometer
59 |  | 10 | Pit 160 17 cm sarake 59 | 17 | 36 | mV | Tensiometer
60 |  | 11 | Pit 160 4 cm sarake 60 | 4 | 36 | mV | Tensiometer
61 |  | 12 | Temp, tensiometer Pit 160 |  |  | °C
62 |  | 13 | Pit 100 31 cm sarake 62 | 31 | 66 | mV | Tensiometer
63 |  | 14 | Pit 100 16 cm sarake 63 | 16 | 36 | mV | Tensiometer
64 |  | 15 | Pit 100 4 cm sarake 64 | 10 | 36 | mV | Tensiometer
65 |  | 16 | T, MASA | Chamber |  | °C | Since 01.06.07

## SunRot
Nokeval 0 | Port 10 and 13 in Mastmet PC
1 | PAR_Tot | 17.01.2017 sensor moved to the measurement tower and logged with tower35m/tow35m_nokevall
2 | PAR_Dif | 17.01.2017 sensor moved to the measurement tower and logged with tower35m/tow35m_nokevall
3 | Sun | 17.01.2017 sensor moved to the measurement tower and logged with tower35m/tow35m_nokevall
4 | Ch4
5 | RH
6 | T_Hydc
7 | T_Pt100
8 | Ch8
9 | Ch9
10 | Ch10
11 | Ch11
12 | Ch12
13 | Ch13
14 | Ch14
15 | Ch15
16 | t_box_Pt

## Soil 04-06-2013
Louhi\D:\smear\soil\SMYYMMDD.maa |  |  |  | Valid since 04.06.2013
Column | Com | Channel | Label | depth of | Length of | Unit | Notes | Conversion
|  |  |  | sensor, cm | tensionmeters
| COM 5 | Noke 0 |  |  | tube / termistor no.
1 |  |  | date and time |  |  | yyyymmddhhmmss
2 |  | 1 | T in cuvette PIT 80 |  |  | °C | Cuvette 13/14/23, since 08.05.2006
3 |  | 2 | T in cuvette PIT 160 |  |  | °C | Cuvette 25
4 |  | 3 | PAR |  |  | mV | Ending 20.11.2007
5 |  | 4 | T in cuvette PIT 70 |  |  | °C | Cuvette 10/20
6 |  | 5 | Pit 50 27 cm | 27 | 66 | mV | Equitensiometer EQ2-358/097
7 |  | 6 | Pit 50 17 cm | 17 | 36 | mV | Equitensiometer EQ2-358/098
8 |  | 7 | Pit 50 4 cm | 4 | 36 | mV | Equitensiometer EQ2-358/099
9 |  | 8 | Pit70 Pt100 temperature |  |  | °C | Pit70 Pt100
10 |  | 9 | Pit 70 16/b cm sarake10 | 16 | 33 | mV | Equitensiometer EQ2-358/096
11 |  | 10 | Pit 70 4 cm sarake 11 | 4 | 33 | mV | Equitensiometer EQ2-358/100
12 |  | 11 | PAR |  |  | mV | Ending 20.11.2007
13 |  | 12 | Empty |  |  |  | T in cuvette PIT 80, until 08.05.2006
14 |  | 13 | T 50 H |  | KTY952 | °C
15 |  | 14 | T 50 A | 5 | KTY953 | °C
16 |  | 15 | T 50 B | 11 | KTY954 | °C
17 |  | 16 | T 50 C | 22 | KTY955 | °C
| COM 5 | Noke 1 | label
18 | 17 | 1 | T 130 H/A | 0 | KTY9517 | °C
19 | 18 | 2 | T 130 A | 5 | KTY9518 | °C
20 | 19 | 3 | T 130 B1 | 9 | KTY9519 | °C
21 | 20 | 4 | T 130 B2 | 23 | KTY9520 | °C
22 | 21 | 5 | T 130 C1 | 58 | KTY9521 | °C
23 | 22 | 6 | T 130 C2 | 90 | KTY9522 | °C
24 | 23 | 7 | T 130 B | 32 | KTY9523 | °C | STONE
25 | 24 | 8 | T 130 C | 102 | KTY9524 | °C | BEDROCK
26 | 25 | 9 | T 160 A | 3 | KTY9526 | °C
27 | 26 | 10 | T 160 B1 | 14 | KTY9527 | °C
28 | 27 | 11 | T 160 B2 | 29 | KTY9528 | °C
29 | 28 | 12 | T 160 H |  | KTY9525 | °C
30 | 29 | 13 | T 160 C1 | 45 | KTY9529 | °C
31 | 30 | 14 | T 160 C2 | 87 | KTY9530 | °C
32 | 31 | 15 | T 160 C3 | 125 | KTY9531 | °C
33 | 32 | 16 | T 160 AIR |  | KTY9532 | °C
| COM 5 | Noke 2 | label
34 | 33 | 1 | PARtst, ending 20.11.2007 |  |  | mV | Since 01.06.09 MaSa chamber open(0), closed(500), closed-sample(900)
35 | 34 | 2 | PARtst |  |  | mV | Ending 30.05.2008
36 | 35 | 3 | PARtst |  |  | mV | Ending 30.05.2008
37 | 36 | 4 | PARtst |  |  | mV | Ending 30.05.2008
38 | 37 | 5 | T in cuvette PIT 70, humus |  | KTY951 | °C | since 051214
39 | 38 | 6 | T 70 H/A | 0 | KTY956 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
40 | 39 | 7 | T 70 B1 | 5 | KTY957 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
41 | 40 | 8 | T 70 B2 | 11 | KTY958 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
42 | 41 | 9 | T 70 B3 | 23 | KTY959 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
43 | 42 | 10 | T 70 C1 | 55 | KTY9510 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
44 | 43 | 11 | T 100 H |  | KTY9511 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
45 | 44 | 12 | T 100 A | 2 | KTY9512 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
46 | 45 | 13 | T 100 B1 | 10 | KTY9513 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
47 | 46 | 14 | T 100 B2 | 22 | KTY9514 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
48 | 47 | 15 | T 100 C1 | 42 | KTY9515 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
49 | 48 | 16 | T 100 C2 | 66 | KTY9516 | °C | -12oC shift in temperature reading between 6. July 2000 and 2. October 2003
| COM 2 | Noke 3 | Label
50 |  | 1 | Theta probe, MASA | 50 |  | mV | Since 01.02.07
51 |  | 2 | Theta probe, MASA | 25 |  | mV | Since 01.02.07
52 |  | 3 | Theta probe, MASA | 10 |  | mV | Since 01.02.07
53 |  | 4 | Theta probe, MASA | H |  | mV | Since 01.02.07
54 |  | 5 | T, MASA | 50 |  | °C | Since 20.10.06
55 |  | 6 | T, MASA | 25 |  | °C | Since 20.10.06
56 |  | 7 | T, MASA | 10 |  | °C | Since 20.10.06
57 |  | 8 | T, MASA | H |  | °C | Since 20.10.06
58 |  | 9 | Pit 160 33 cm sarake 58 | 33 | 66 | mV | Equitensiometer EQ2-359/001
59 |  | 10 | Pit 160 17 cm sarake 59 | 17 | 36 | mV | Equitensiometer EQ2-359/002
60 |  | 11 | Pit 160 4 cm sarake 60 | 4 | 36 | mV | Equitensiometer EQ2-359/006
61 |  | 12 | Pit100 Pt100 temperature |  |  | °C | Pit100 Pt100
62 |  | 13 | Pit 100 31 cm sarake 62 | 31 | 66 | mV | Equitensiometer EQ2-359/004
63 |  | 14 | Pit 100 16 cm sarake 63 | 16 | 36 | mV | Equitensiometer EQ2-359/005
64 |  | 15 | Pit 100 4 cm sarake 64 | 10 | 36 | mV | Equitensiometer EQ2-359/007
65 |  | 16 | T, MASA chamber Pt100 | Chamber |  | °C | Since 01.06.07, Pt100

## MaSa
Louhi\D:\smear\MaSa\SMYYMMDD.msa |  |  |  | Began 01.06.2009
Program: SMEAR MaSa
Datafile column | Nokeval 2 Channel | Measured Quantity | Depth [cm] | Unit in datafile | Notes
1 |  | date and time |  | yyyymmddhhmmss | Nearby Pit50
2 | 1 | Chamber open/closed/closed-sample |  | 0/495/900 mV
3 | 2 | Empty
4 | 3 | Empty
5 | 4 | Empty
6 | 5 | Empty
7 | 6 | Empty
8 | 7 | Empty
9 | 8 | Empty
10 | 9 | Empty
11 | 10 | Empty
12 | 11 | Empty
13 | 12 | Empty
14 | 13 | Empty
15 | 14 | Empty
16 | 15 | Empty
17 | 16 | Empty
Nokeval 3 Channel
18 | 1 | Theta probe, MASA | 50 | mV
19 | 2 | Theta probe, MASA | 25 | mV
20 | 3 | Theta probe, MASA | 10 | mV
21 | 4 | Theta probe, MASA | H | mV
22 | 5 | T, MASA | 50 | °C
23 | 6 | T, MASA | 25 | °C
24 | 7 | T, MASA | 10 | °C
25 | 8 | T, MASA | H | °C
26 | 9 | Empty
27 | 10 | Empty
28 | 11 | Empty
29 | 12 | Empty
30 | 13 | Empty
31 | 14 | Empty
32 | 15 | Empty
33 | 16 | T, MASA | Chamber | °C

## SoilCH
Louhi\D:\smear\vkammio\SMYYMMDD.lis |  |  | Began 23.4.2008
Program: SOILCH
Datafile column | Nokeval 8 Channel | Measured Quantity | Unit in datafile | Notes
1 |  | date and time | yyyymmddhhmmss | Nearby Pit100
2 | 7 | chamber open/closed | 1137/680 mV
3 | 8 | Temperature in chamber | °C
4 | 9 | CO2 | ppm
5 | 10 | LiCor PAR | mV
6 | 11 | Empty
7 | 12 | Empty

## ToHeLi_1-2
Louhi\D:\smear\toheli\SMYYMMDD.li1 |  |  | Began 22.4.2009
Program: ToHeLi_1-2
Datafile column | Nokeval 8 Channel | Measured Quantity | Unit in datafile | Notes
1 |  | date and time | yyyymmddhhmmss | Nearby Pit100
2 | 15 | Chamber1 CO2 | ppm
3 | 16 | Chamber1 open/closed | 270/4.8 mV
Nokeval 9 Channel
4 | 1 | Chamber1 T air | °C
5 | 2 | Chamber1 T ground | °C
6 | 3 | Chamber1 PAR 1 | mV
7 | 4 | Chamber1 PAR 2 | mV
8 | 5 | Chamber1 RH | mV
9 | 6 | Chamber1 T (RH)
10 | 7 | Chamber2 CO2 | ppm
11 | 8 | Chamber2 open/closed | 1222/5 mV
12 | 9 | Chamber2 T air | °C
13 | 10 | Chamber2 T ground | °C
14 | 11 | Chamber2 PAR 1 | mV
15 | 12 | Chamber2 PAR 2 | mV
16 | 13 | Chamber2 RH | mV
17 | 14 | Chamber2 T (RH)
Since 24.04.2015 11:30
Louhi\D:\smear\toheli\SMYYMMDD.li1
Program: ToHeLi_1-2
Datafile column | Nokeval 8 Channel | Measured Quantity | Unit in datafile
1 |  | date and time | yyyymmddhhmmss
2 | 5 | Chamber 1 air
3 | 6 | Chamber 1 soil
4 | 15 | Chamber1 CO2 | ppm
5 | 16 | Chamber1 open/closed | 270/4.8 mV
Nokeval 9 Channel
5 | 1 | Chamber1 T air | °C
6 | 2 | Chamber1 T ground | °C
7 | 3 | Chamber1 PAR 1 | mV
8 | 4 | Chamber1 PAR 2 | mV
9 | 5 | Chamber1 RH | mV
10 | 6 | Chamber1 T (RH)
11 | 7 | Chamber2 CO2 | ppm
12 | 8 | Chamber2 open/closed | 1222/5 mV
13 | 9 | Chamber2 T air | °C
14 | 10 | Chamber2 T ground | °C
15 | 11 | Chamber2 PAR 1 | mV
16 | 12 | Chamber2 PAR 2 | mV
17 | 13 | Chamber2 RH | mV
18 | 14 | Chamber2 T (RH)

## ToHeLi_3
Sampo\D:\smear\toheli3\SMYYMMDD.li3 |  |  | Began 18.5.2009, modified 17.06.2009
Program: ToHeLi_3
Datafile column | Nokeval 1 Channel | Measured Quantity | Unit in datafile | Notes
1 |  | date and time | yyyymmddhhmmss | Nearby Pit80
2 | 9 | Chamber3 CO2 | ppm
3 | 10 | Chamber3 open/closed | 272/5.9 mV
4 | 11 | Chamber3 T air | °C
5 | 12 | Chamber3 T ground | °C
6 | 13 | Chamber3 PAR 1 | mV
7 | 14 | Chamber3 PAR 2 | mV
8 | 15 | Chamber3 RH | mV
9 | 16 | Chamber3 T (RH) | mV

## Eliza2
Sampo\D:\smear\v2kammio\SMYYMMDD.V2K |  |  | Between 29.5.2008 - 09.11.2008
Program: Eliza2
Datafile column | Nokeval 1 Channel | Measured Quantity | Unit in datafile | Notes
1 |  | date and time | yyyymmddhhmmss | Nearby Pit80
2 | 1 | CO2 in chamber | ppm
3 | 4 | Chamber open/closed | 1422 / 81 mV
4 | 13 | T1 in chamber | °C
5 | 14 | T2 | °C

## Sun
SUN |  | smyymmdd.sun
Nokeval configuration | Nokeval 0, Port 6 in Pepe PC | Updated 12.01.2000 |  |  | notation in Dataviewer
Datafile column | channel | label | Description | units | SUN | Notes
1 |  | date and time |  | yyyymmhhss
2 | 1 | UVA | UVA radiation intensity | mV | UVA
3 | 2 | UVB | UVB radiation intensity | mV | UVB
4 | 3 | T UVA | Sensor temperature | mV | UVAT
5 | 4 | T UVB | Sensor temperature | mV | UVBT
6 | 5 | Empty
7 | 6 | Empty
8 | 7 | RAIN | Rain amount | mV | RAIN | Tipping bucket position
9 | 8 | SWS | Rain detector | mV | SWS
10 | 9 | Direct PAR | Direct PAR intensity | mV | SPAR ? | Sensor unconnected
11 | 10 | Diffuse | Diffuse radiation | mV | DR ?
12 | 11 | Direct | Direct global radiation intensity | mV | SR ? | Sensor unconnected
13 | 12 | NET | Net radiation intensity | mV | NET | Sensor unconnected
14 | 13 | Reflected | Reflected global radiation intensity | mV | RK | Sensor unconnected
15 | 14 | Global | Global radiation intensity | mV | GR
16 | 15 | Refl. PAR | Reflected PAR intensity | mV | RPAR | Sensor unconnected
17 | 16 | PAR | PAR intensity | mV | PAR
SMEAR II |  | Column order in Sun-datafiles (*.sun)
Valid 15.03.2000 - (see Notes for some details)
Louhi\D:\smear\sun\SMYYMMDD.SUN
Datafile column | Nokeval 12, Port 5 in Louhi PC | Measured Quantity | Unit in datafile | Description | Notes
1 | channel | date & time | yyyymmddhhmmss
2 | 1 | UVA | mV | UVA radiation | signal+unit changed 17.06.1997
3 | 2 | UVB | mV | UVB radiation | signal+unit changed 17.06.1997
4 | 3 | T UVA | mV | sensor temperature | signal+unit changed 17.06.1997
5 | 4 | T UVB | mV | sensor temperature | signal+unit changed 17.06.1997
6 | 5 | Empty
7 | 6 | Empty
8 | 7 | Rain | mV | Rain amount | tipping bucket position
9 | 8 | SWS | mV | Rain detector
10 | 9 | Empty
11 | 10 | Diffuse | mV | diffuse radiation | sensor + shadow ring installed 15.03.2000
12 | 11 | Empty
13 | 12 | Empty
14 | 13 | Empty
15 | 14 | Global | mV | global radiation | signal+unit changed 17.06.1997
16 | 15 | Empty
17 | 16 | PAR | mV | PAR radiation | signal+unit changed 17.06.1997
SMEAR II |  | Column order in Sun-datafiles (*.sun)
Valid 22.04.2009 - (see Notes for some details)
Louhi\D:\smear\sun\SMYYMMDD.SUN
Datafile column | Nokeval 12, Port 5 in Louhi PC | Measured Quantity | Unit in datafile | Description | Notes
1 | channel | date & time | yyyymmddhhmmss
2 | 1 | UVA | mV | UVA radiation | signal+unit changed 17.06.1997
3 | 2 | UVB | mV | UVB radiation | signal+unit changed 17.06.1997
4 | 3 | T UVA | mV | sensor temperature | signal+unit changed 17.06.1997
5 | 4 | T UVB | mV | sensor temperature | signal+unit changed 17.06.1997
6 | 5 | PRI up1 |  |  | Began 22.04.2009
7 | 6 | PRI up2 |  |  | Began 22.04.2009
8 | 7 | Rain | mV | Rain amount | tipping bucket position
9 | 8 | SWS | mV | Rain detector
10 | 9 | Empty
11 | 10 | Diffuse S/N 0895 | mV | diffuse radiation | sensor + shadow ring installed 15.03.2000
12 | 11 | PRI down1 |  |  | Began 22.04.2009
13 | 12 | PRI down2 |  |  | Began 22.04.2009
14 | 13 | Empty
15 | 14 | Global | mV | global radiation | signal+unit changed 17.06.1997
16 | 15 | Empty
17 | 16 | PAR | mV | PAR radiation | signal+unit changed 17.06.1997
SMEAR II |  | Column order in Sun-datafiles (*.sun)
Valid 07.12.2009 - (see Notes for some details)
Louhi\D:\smear\sun\SMYYMMDD.SUN
Datafile column | Nokeval 12, Port 5 in Louhi PC | Measured Quantity | Unit in datafile | Description | Notes
1 | channel | date & time | yyyymmddhhmmss
2 | 1 | UVA | mV | UVA radiation | signal+unit changed 17.06.1997 | 17.01.2017 sensor moved to the measurement tower and logged with tower35m/tow35m_nokevall
3 | 2 | UVB | mV | UVB radiation | signal+unit changed 17.06.1997 | 17.01.2017 sensor moved to the measurement tower and logged with tower35m/tow35m_nokevall
4 | 3 | T UVA | mV | sensor temperature | signal+unit changed 17.06.1997 | 17.01.2017 sensor moved to the measurement tower and logged with tower35m/tow35m_nokevall
5 | 4 | T UVB | mV | sensor temperature | signal+unit changed 17.06.1997 | 17.01.2017 sensor moved to the measurement tower and logged with tower35m/tow35m_nokevall
6 | 5 | PAR total | µmol/m2s | Delta Sunshine | Began 07.12.2009 | Broken 15.11.2014, new sensor logged with sunrot program
7 | 6 | PAR diffuse | µmol/m2s | Delta Sunshine | Began 07.12.2009 | Broken 15.11.2014, new sensor logged with sunrot program
8 | 7 | Rain | mV | Rain amount | tipping bucket position | Measurement stopped 17.01.2017
9 | 8 | SWS | mV | Rain detector |  | Measurement stopped 17.01.2018
10 | 9 | Empty
11 | 10 | Diffuse S/N 0895 | mV | diffuse radiation | sensor + shadow ring installed 15.03.2000 and removed 18.03.2010
12 | 11 | PRI down1 |  |  | Began 22.04.2009, end 05.10.2009
13 | 12 | PRI down2 |  |  | Began 22.04.2009, end 05.10.2009
14 | 13 | Empty
15 | 14 | Global | mV | global radiation | signal+unit changed 17.06.1997 | 17.01.2017 sensor moved to the measurement tower and logged with tower35m/tow35m_nokevall
16 | 15 | Empty
17 | 16 | PAR | mV | PAR radiation | signal+unit changed 17.06.1997 | 17.01.2017 sensor moved to the measurement tower and logged with tower35m/tow35m_nokevall

## LWS
Louhi\D:\smear\lws\SMYYMMDD.tul |  |  |  | 3.6. - 3.10.2005
Datafile column | Nokeval Channel | Measured Quantity | Unit in datafile | Description | Notes
1 |  | date & time | yyyymmddhhmmss
2 | 1 | Empty
3 | 2 | Empty
4 | 3 | Empty
5 | 4 | Empty
6 | 5 | Empty
7 | 6 | Empty
8 | 7 | Empty
9 | 8 | Empty
10 | 9 | Empty
11 | 10 | Empty
12 | 11 | LWS | mV ? | Leaf Wetness Sensor | No: 6
13 | 12 | LWS | mV ? | Leaf Wetness Sensor | No: 5
14 | 13 | LWS | mV ? | Leaf Wetness Sensor | No: 4
15 | 14 | LWS | mV ? | Leaf Wetness Sensor | No: 3
16 | 15 | LWS | mV ? | Leaf Wetness Sensor | No: 2
17 | 16 | LWS | mV ? | Leaf Wetness Sensor | No: 1
Louhi\D:\smear\lws\SMYYMMDD.tul |  |  |  | Began 12.05.2007
Datafile column | Nokeval 1 Channel | Measured Quantity | Unit in datafile
1 |  | date & time | yyyymmddhhmmss
2 | 11 | LWS | mV
3 | 12 | LWS | mV
4 | 13 | LWS | mV
5 | 14 | LWS | mV
6 | 15 | LWS | mV
7 | 16 | LWS | mV

## Tree 05.06.96
Smyymmdd.puu |  | 05.06.1996 - 30.05.2001
Column | Source | Meaning | Units | Range of values | Example | Liitin, boxi | Wire colour
1 |  | Date and time | yyyymmddhhmm |  | 199706050830
2 | Nok 9, Ch1 | Sylvac, Hanna, stem, lowest | 1e-4 m | 0 - 650 | 246.15
3 | Nok 9, Ch2 | Sylvac, Hanna, stem, 2nd lowest | 1e-4 m | 0 - 650 | 246.15
4 | Nok 9, Ch3 | Sylvac, Hanna, stem, 2nd highest, removed | 1e-4 m | 0 - 650 | 246.15
5 | Nok 9, Ch4 | Sylvac, Hanna, stem, highest | 1e-4 m | 0 - 650 | 246.15
6 | Nok 9, Ch5 | Sylvac, Hanna, branch | 1e-4 m | 0 - 650 | 246.15
7 | Nok 9, Ch6 | Sylvac, Laura, crown base | 1e-4 m | 0 - 650 | 246.15
8 | Nok 9, Ch7 | Sylvac, Laura | 1e-4 m | 0 - 650 | 246.15
9 | Nok 9, Ch8 | Sylvac, | 1e-4 m | 0 - 650 | 246.15
10 | Nok 9, Ch9 | Thermocouple Hanna 2 | °C
11 | Nok 9, Ch10 | Thermocouple Hanna, next to ELPA top 30.6.98 | °C
12 | Nok 9, Ch11 | ELPA input voltage
13 | Nok 9, Ch12 | Thermocouple ? | °C
14 | Nok 9, Ch13 | Thermocouple Taru stem resp. chamber 30.6.98 | °C
15 | Nok 9, Ch14 | Thermocouple Hanna 1 | °C
16 | Nok 9, Ch15 | ELPA, Hanna, top
17 | Nok 9, Ch16 | ELPA, Hanna, base
18 | Nok 4, Ch 1 | Tensionmeters
19 | Nok 4, Ch 2 | Tensionmeters
20 | Nok 4, Ch 3 | Tensionmeters
21 | Nok 4, Ch 4 | Tensionmeters
22 | Nok 4, Ch 5 | Tensionmeters
23 | Nok 4, Ch 6 | Tensionmeters NEGATIVE SIGNAL
24 | Nok 4, Ch 7 | Tensionmeters
25 | Nok 4, Ch 8 | Tensionmeters
26 | Nok 4, Ch 9 | Tensionmeters
27 | Nok 4, Ch 10 | Tensionmeters
28 | Nok 4, Ch 11 | Tensionmeters
29 | Nok 4, Ch 12 | Tensionmeters
30 | Nok 4, Ch 13 | Tensionmeters
31 | Nok 4, Ch 14 | Tensionmeters NEGATIVE SIGNAL
32 | Nok 4, Ch 15 | Tensionmeters
33 | Nok 4, Ch 16 | Tensionmeters NEGATIVE SIGNAL
34 | Nok 5, Ch 1 | Kfiraos, empty | mV
35 | Nok 5, Ch 2 | Kfiraos, empty | mV
36 | Nok 5, Ch 3 | Kfiraos Input Voltage | mV
37 | Nok 5, Ch 4 | Kfiraos 1, Laura, top | mV |  |  | 1 lower | green, yellow
38 | Nok 5, Ch 5 | Kfiraos 2, Kaisa, branch | mV |  |  | 2 lower | white, black
39 | Nok 5, Ch 6 | Kfiraos 5, Tiina, top | mV |  |  | 3 lower | violet, brown
40 | Nok 5, Ch 7 | Kfiraos 4, Nuria, top | mV |  |  | 4 lower | orange, pink
41 | Nok 5, Ch 8 | Kfiraos 3, broken, not installed | mV |  |  | 5 lower | turq., grey
42 | Nok 5, Ch 9 | Kfiraos 6, Eija, top | mV |  |  | 1 upper | green, yellow
43 | Nok 5, Ch 10 | Kfiraos 7, Laura, branch | mV |  |  | 2 upper | white, black
44 | Nok 5, Ch 11 | Kfiraos 8, Eija branch | mV |  |  | 3 upper | violet, brown
45 | Nok 5, Ch 12 | empty | mV |  |  | 4 upper | orange, pink
46 | Nok 5, Ch 13 | empty | mV |  |  | 5 upper | turq., grey
47 | Nok 5, Ch 14 | empty | mV
48 | Nok 5, Ch 15 | empty | mV
49 | Nok 5, Ch 16 | empty | mV

## Tree 30.05.01
Smyymmdd.puu |  | Configuration starting at 2001/05/30 |  |  | Configuration starting at 2002/05/16
Column | Source | Meaning | Column | Source | Meaning | Notes
1 |  | Date and time | 1 |  | Date and time
2 | Nok 9, Ch1 | Solartron 1 | 2 | Nok 9, Ch1 | Solartron 1
3 | Nok 9, Ch2 | Solartron 2 | 3 | Nok 9, Ch2 | Solartron 2
4 | Nok 9, Ch3 | Solartron 3 | 4 | Nok 9, Ch3 | Solartron 3
5 | Nok 9, Ch4 | Solartron 4 | 5 | Nok 9, Ch4 | Solartron 4
6 | Nok 9, Ch5 | Solartron 5 | 6 | Nok 9, Ch5 | Solartron 5
7 | Nok 9, Ch6 | Solartron 6 | 7 | Nok 9, Ch6 | Solartron 6
8 | Nok 9, Ch7 | Solartron 7 | 8 | Nok 9, Ch7 | Solartron 7
9 | Nok 9, Ch8 | Solartron 8 | 9 | Nok 9, Ch8 | Solartron 8
10 | Nok 9, Ch9 | Temp. For Solartron 1 | 10 | Nok 9, Ch9 | Temp. For Solartron 1
11 | Nok 9, Ch10 | Temp. For Solartron 2 | 11 | Nok 9, Ch10 | Temp. For Solartron 2
12 | Nok 9, Ch11 | Temp. For Solartron 3 | 12 | Nok 9, Ch11 | Temp. For Solartron 3
13 | Nok 9, Ch12 | Temp. For Solartron 4 | 13 | Nok 9, Ch12 | Temp. For Solartron 4
14 | Nok 9, Ch13 | Temp. For Solartron 5 | 14 | Nok 9, Ch13 | Temp. For Solartron 5
15 | Nok 9, Ch14 | Temp. For Solartron 6 | 15 | Nok 9, Ch14 | Temp. For Solartron 6
16 | Nok 9, Ch15 | Temp. For Solartron 7 | 16 | Nok 9, Ch15 | Temp. For Solartron 7
17 | Nok 9, Ch16 | Temp. For Solartron 8 | 17 | Nok 9, Ch16 | Temp. For Solartron 8
18 | Nok 5, Ch 1 | LWS 1 | 18 | Nok 5, Ch 1 | LWS 1, Unit1 (south) | Sensor is empty
19 | Nok 5, Ch 2 | LWS 2 | 19 | Nok 5, Ch 2 | LWS 2, Unit1 (south) | Tree Hanna
20 | Nok 5, Ch 3 | LWS 3 | 20 | Nok 5, Ch 3 | LWS 3, Unit1 (south) | Tree Hanna
21 | Nok 5, Ch 4 | LWS 4 | 21 | Nok 5, Ch 4 | LWS 6, Unit1 (south) | Inside cuvette 118
22 | Nok 5, Ch 5 | LWS 5 | 22 | Nok 5, Ch 5 | LWS 5, Unit1 (south) | Tree Hanna
23 | Nok 5, Ch 6 | LWS 6 | 23 | Nok 5, Ch 6 | LWS 4, Unit1 (south) | Tree Hanna
24 | Nok 5, Ch 7 | Granier 1 | 24 | Nok 5, Ch 7 | Granier 1
25 | Nok 5, Ch 8 | Granier 2 | 25 | Nok 5, Ch 8 | Granier 2
26 | Nok 5, Ch 9 | Granier 3 | 26 | Nok 5, Ch 9 | LWS 1, Unit2 (north) | Sensor is empty
27 | Nok 5, Ch 10 | Solartron 10 | 27 | Nok 5, Ch 10 | LWS 2, Unit2 (north) | Tree Liisa
28 | Nok 5, Ch 11 | Solartron 11 | 28 | Nok 5, Ch 11 | LWS 3, Unit2 (north) | Tree Liisa
29 | Nok 5, Ch 12 | Solartron 12 | 29 | Nok 5, Ch 12 | LWS 4, Unit2 (north) | Inside cuvette 111
30 | Nok 5, Ch 13 | Not connected | 30 | Nok 5, Ch 13 | LWS 5, Unit2 (north) | Tree Liisa
31 | Nok 5, Ch 14 | Not connected | 31 | Nok 5, Ch 14 | LWS 6, Unit2 (north) | Tree Liisa
32 | Nok 5, Ch 15 | Not connected | 32 | Nok 5, Ch 15 | LWS 7, Unit2 (north) | Tree Liisa
33 | Nok 5, Ch 16 | Not connected | 33 | Nok 5, Ch 16 | LWS 8, Unit2 (north) | Tree Liisa

## Tree 07.10.03
Smyymmdd.puu |  |  |  | Configuration starting at 2003/10/07 |  |  | Last modified at 16.2.2005
Last modified at 17.2.2005
Column | Box addr | Sensor type | Tree | SensorId / Place | Height | Date installed | PIC | Date changed | PIC | Date removed | PIC | Comment
1 |  |  |  | Date and time
2 | E1; Nok3,1 | Solar | 1 | T1BPh | 1.42 | 7.10.2003 | SS&MP&EN&TH |  |  | 2006-05-31 |  | Huom! Piuhat ristissä 1:ssä ja 2:ssa
3 | E2; Nok3,2 | Solar | 1 | T1BX | 1.42 | 7.10.2003 | SS&MP&EN&TH |  |  | 2006-05-31 |  | Huom! Piuhat ristissä 1:ssä ja 2:ssa
4 | E3; Nok3,3 | Solar | 1 | T1TX | 12.72 | 7.10.2003 | SS&MP&EN&TH |  |  | 2006-05-31
5 | E4; Nok3,4 | Solar | 1 | T1TPh | 12.72 | 7.10.2003 | SS&MP&EN&TH |  |  | 2006-05-31
6 | E5; Nok3,5 | Solar | 2 | T2BX | 2.4 | 7.10.2003 | SS&MP&EN&TH
7 | E6; Nok3,6 | Solar | 2 | T2BPh | 2.4 | 7.10.2003 | SS&MP&EN&TH
8 | E7; Nok3,7 | Solar | 2 | T2TX | 10.65 | 7.10.2003 | SS&MP&EN&TH
9 | E8; Nok3,8 | Solar | 2 | T2TPh | 10.65 | 7.10.2003 | SS&MP&EN&TH
10 | E9; Nok3,9 | Solar | 3 | T3BX | 2.1 | 7.10.2003 | SS&MP&EN&TH
11 | E10; Nok3,10 | Solar | 3 | T3BPh | 2.1 | 7.10.2003 | SS&MP&EN&TH
12 | E11; Nok3,11 | Solar | 3 | T4TX | 9.87 | 7.10.2003 | SS&MP&EN&TH
13 | E12; Nok3,12 | Solar | 3 | T3TPh | 9.87 | 7.10.2003 | SS&MP&EN&TH
14 | Nok3, 13 | TC |  | T ?
15 | Nok3, 14 | TC | 2 | T2B | 2.4 | syksy 2003 | SS&MP&EN&TH
16 | Nok3, 15 | TC | 1 | T1B | 1.42 | syksy 2003 | SS&MP&EN&TH
17 | Nok3, 16 | TC | 3 | T3B | 2.1 | syksy 2003 | SS&MP&EN&TH
18 | Nok1, 1 | TC | 1 | T1T | 12.72 | syksy 2003 | SS&MP&EN&TH | 15.2.2005 | ES&MP |  |  | Korjattu Nokevallin konfigurointi, data tähän asti puppua
19 | Nok1, 2 | TC | 2 | T2T | 10.65 | syksy 2003 | SS&MP&EN&TH | 15.2.2006 | ES&MP |  |  | Korjattu Nokevallin konfigurointi, data tähän asti puppua
20 | Nok1, 3 | TC | 3 | T3T | 9.87 | syksy 2003 | SS&MP&EN&TH | 15.2.2007 | ES&MP |  |  | Korjattu Nokevallin konfigurointi, data tähän asti puppua
21 | Nok1, 4 | TC | 1 | verso,ylä | 14 | 16.2.2005 | TH
22 | Nok1, 5 | TC | 1 | verso,anturitasolla | 12.72 | 16.2.2005 | TH
23 | Nok1, 6 | TC | 1 | verso,ylä | 14 | 16.2.2005 | TH
24 | Nok1, 7 | TC | 1 | runko, anturitaso | 12.72 | 16.2.2005 | TH
25 | Nok1, 8 | TC | 1 | oksa, ylä | 14 | 16.2.2005 | TH
26 | Nok1, 9 | TC | 1 | oksa, anturitaso | 12.72 | 16.2.2005 | TH
27 | Nok1, 10 | TC |  | ???
28 | Nok1, 11 | LWS |  | Tree1 |  | 2007-05-12
29 | Nok1, 12 | LWS |  | Tree1 |  | 2007-05-12
30 | Nok1, 13 | LWS |  | Tree1 |  | 2007-05-12
31 | Nok1, 14 | LWS |  | Tree1 |  | 2007-05-12
32 | Nok1, 15 | LWS |  | Tree1 |  | 2007-05-12
33 | Nok1, 16 | LWS |  | Tree1 |  | 2007-05-12
Tree sensors | Solar = Solartron
TC= thermocouple
Granier=Sap flow needle pair
Trees | 1 | Twin-headed,front of the tower
| 2 | Thicker, right, behind the tower
| 3 | Thinner, left, behind the tower
[5 empty] | Tree3 |  | Tree2
Tree1
SensorId/Place | TnTX=Tree n, Top, Xylem
Boxes: | E= Erckens white box
Noknn = Nokeval Module nn
Dyn = Dynamax

## KFIRAOS
KFIRAOS INSTALLATION |  |  | 980518
| nokeval | tree | computer
| channel |  | column
kfiraos 1 | 4 | Laura, top | 36
kfiraos 2 | 5 | Kaisa, branch | 37
kfiraos 5 | 6 | Tiina, top | 38
kfiraos 4 | 7 | Nuria, top | 39
kfiraos 6 | 9 | Eija, top | 41
kfiraos 7 | 10 | Laura, branch | 42
kfiraos 8 | 11 | Not installed | 43
kfiraos 3 |  | Not installed, broken

## SHARP
SHARPin tulostusformaatiksi valittu 7 eli tiedostossa rivit:
rivinvaihdon tulkinnasssa on sanomista, muuten pelaa...
(ekin timestamp), =20
date,time,status,concentration,mass,flowrate,RH%,T1,T2,T3,T4,heatercycle%,P1=
_hPa,P2_hPa,P3_hPa,betaconc-short,betaconc-long,nephconc,nephconc-long,neph-=
calfactor,betacounts
noista varmaan t=C3=A4rkeimm=C3=A4t katteluohjelmaan on,
status,concentration,mass,flowrate
T1 ambient temperature sensor (katolla inletin vieress=C3=A4)
T2 sample temperature sensor (n=C3=A4ytekammiossa, filtterinauha)
T3 orifice temperature sensor (virtausmittauskuristuksen l=C3=A4mp=C3=B6tila=
, =20
virtausmittaus siis paine-eroanturilla)
T4 heater outer wall temperature sensor (inletputken l=C3=A4mp=C3=B6tila)
P1 orifice differential pressure (virtaus raakasignaali)
P2 differential vacuum under filter tape (filtterinauhan aiheuttama =20
paineh=C3=A4vi=C3=B6)
P3 barometric pressure
RH sample relative humidity

## Envitems
Time | Temperature | Relative Humidity | Barometric Pressure | NO2 | CO | O3 | NO | PM2.5 | PM10 | Uptime